## Metadata
cruise |  | 2008-27
ship |  | Tully
dates |  | August 13-August 28 2008
Chief Scientist |  | Marie Robert
Reference: | Miller, L.A., M. Davelaar, W.K. Johnson, and J. Linguanti, 2009. Line P Inorganic Carbon Data Set, Cruise 2008-27, August 2008. Institute of Ocean Sciences, Fisheries and Oceans Canada, Sidney, BC, Canada.
TIC and alkalinity collected in 500-ml glass bottles with ground glass stoppers sealed with Apiezon M grease, elastic bands, and plastic clips.
Poisoned w/ 100 microlitres saturated HgCl2
Alkalinity and TIC analyses were stored at 4 C in the dark and analyzed ashore from the same glass bottle.
For raw data and details on CTD and nutrient data, see original data archived at the Institute of Ocean Sciences, Sidney, BC, Canada
flags
2 = good samle
3 = suspect sample
4 = bad sample
6 = mean of duplicate samples
9 = missing sample
Total Inorganic Carbon
Coulometric titration
| Reference: |  | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.
technique for standardization |  |  | CRM water
analysis temperature (C) |  |  | 20
Standard used:
Dickson CRM |  |  | 86
certified value |  |  | 1998.37+/-0.54
IOS standard |  |  | 14
calibrated value |  |  | 2036.68+/-0.45
correction formula
(known standard value/measured standard value) X analyzed sample value X 1.0002 (HgCl2 correction)
Notes: | 32 | Unstable analysis
| 35 | Low relative to historical experience, inconsistent w/alk & pH
| 44 | Inconsistent w/Alk & pH
| 211 | C parameters inconsistent
| 555 | Large difference between replicates
| 557 | Different from replicate niskins
Replicate Analyses (same niskin)
Sample | A | B |  | Difference | Absolute
id | umol/kg | umol/kg | average | umol/kg | Difference
28 | 2381.8747 | 2381.699 | 2381.699 | -0.17572 | 0.17572
193 | 2366.5437 | 2367.2133 | 2366.8785 | 0.66959 | 0.66959
310 | 2360.1556 | 2360.6468 | 2360.4012 | 0.49114 | 0.49114
364 | 2367.5659 | 2367.1035 | 2367.3347 | -0.46244 | 0.46244
429 | 2354.0389 | 2354.2769 | 2354.1579 | 0.23802 | 0.23802
553 | 2390.6005 | 2390.3156 | 2390.4581 | -0.28484 | 0.28484
554 | 2391.1293 | 2390.9616 | 2391.0455 | -0.16764 | 0.16764
555 | 2391.4804 | 2390.5825 | 2391.0315 | -0.89793 | 0.89793
556 | 2390.8271 | 2391.4783 | 2391.1527 | 0.65119 | 0.65119
557 | 2386.9025 | 2387.0905 | 2386.9965 | 0.18799 | 0.18799
average variability: |  |  |  | 0.024935 | 0.42265
standard deviation of variability: |  |  |  | 0.51204 | 0.25376
Replicate Samples (different niskins)
Sample |  | Average | STDDEV. | difference | absolute
id | umol/kg | umol/kg | umol/kg | umol/kg | difference
27 | 2382.3591
28 | 2381.699
|  | 2382.0291 |  | -0.66014 | 0.66014
192 | 2367.4509
193 | 2366.8785
|  | 2367.1647 |  | -0.57242 | 0.57242
310 | 2360.4012
311 | 2359.547
|  | 2359.9741 |  | -0.8542 | 0.8542
364 | 2367.3347
365 | 2367.7994
|  | 2367.567 |  | 0.46473 | 0.46473
553 | 2390.4581
554 | 2391.0455
555 | 2391.0315
556 | 2391.1527
557 | 2386.9965
|  | 2390.1368 | 1.7764
average variability: |  |  |  | -0.40551 |  | 0.86559
standard deviation of variability: |  |  |  | 0.59199 |  | 0.52883
Precision estimate: |  | 0.87 | umol/kg
Total Alkalinity
Potentiometric titration
| Reference: |  | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.
Cell type |  |  | Open
Endpoint determination |  |  | Modified gran plot
Standard used:
Dickson CRM |  |  | 86
certified value |  |  | 2175.56 ± 0.67
Dickson CRM |  |  | 88
certified value |  |  | 2205.41 ± 0.43
analysis temperature (C) |  |  | 20
Acid batch |  |  | 6
acid concentration |  |  | 0.098069
sample weight |  |  | 70-80 grams
correction formula
(known standard value/measured standard value) X analyzed sample value X 1.0002 (HgCl2 correction)
Notes: | 29 | Poor calibration
| 30 | High outlier, inconsistent w/TIC & pH
| 44 | Inconsistent w/TIC & pH
| 192 | Poor calibration
| 202 | High outlier; inconsistent w/TIC & pH
| 209 | Poor calibration
| 210 | Poor calibration
| 211 | C parameters inconsistent
| 314 | Poor endpoint; inconsistent w/TIC & pH
| 318 | C parameters inconsistent; Poor endpoint
| 321 | High outlier, unstable analysis
| 383 | Poor endpoint; inconsistent w/TIC & pH
| 445 | High outlier
| 445 | High outlier
| 502 | Poor calibration
| 504 | Poor calibration; inconsistent w/TIC & pH
| 507 | Poor calibration
| 508 | Poor calibration
Replicate Analyses (same niskin)
Sample | A | B | Average | Difference | Absolute
id | umol/kg | umol/kg | umol/kg | umol/kg | Difference
28 | 2395.1979 | 2394.8776 | 2395.0378 | 0.32025 | 0.32025
193 | 2430.2055 | 2429.194 | 2430.2055 | 1.0115 |  | Poor calibration
310 | 2430.3091 | 2430.9268 | 2430.6179 | -0.61768 | 0.61768
364 | 2429.0085 | 2428.4182 | 2428.7133 | 0.59034 | 0.59034
429 | 2425.7016 | 2424.5344 | 2425.118 | 1.1672 | 1.1672
553 | 2420.8468 | 2421.1995 | 2421.0231 | -0.35274 | 0.35274
554 | 2421.1984 | 2420.7191 | 2420.9587 | 0.47928 | 0.47928
555 | 2420.4201 | 2421.2611 | 2420.8406 | -0.84105 | 0.84105
556 | 2419.2004 | 2419.0183 | 2419.1094 | 0.1821 | 0.1821
557 | 2420.36 | 2420.813 | 2420.5865 | -0.45294 | 0.45294
average variability: |  |  |  | 0.14862 | 0.55595
standard deviation of variability: |  |  |  | 0.69097 | 0.29854
Replicate Samples (different niskins)
Sample |  | Average | STDDEV. | difference | absolute
id | umol/kg | umol/kg | umol/kg | umol/kg | difference
27 | 2394.047
28 | 2395.0378
|  | 2394.5424 |  | 0.99078 | 0.99078
192 | 2428.5843 |  |  |  |  | Poor calibration
193 | 2430.2055
|  | 2429.3949 |  | 1.6212
310 | 2430.6179
311 | 2430.6092
|  | 2430.6136 |  | -0.0087072 | 0.0087072
364 | 2428.7133
365 | 2428.4864
|  | 2428.5998 |  | -0.22696 | 0.22696
553 | 2421.0231
554 | 2420.9587
555 | 2420.8406
556 | 2419.1094
557 | 2420.5865
|  | 2420.5037 | 0.79707
average variability: |  |  |  | 0.59408 |  | 0.50588
standard deviation of variability: |  |  |  | 0.86599 |  | 0.46364
Precision estimate: |  | 0.56 | umol/kg
pH
All samples were collected directly from the rosette into 10-cm spectrophotometer cells
Spectrophotometric
| Reference: |  | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.
pH scale: |  | Total hydrogen ion concentration
Analysis temperature: |  | 25 C
Dye: |  |  |  |  | m-Cresol Purple
| Concentration: |  |  |  | ≥ 2 mmol/L
| pH: |  |  |  | 7.96
| Absorbance ratio (578nm/474nm): |  |  |  | 1.47
| pH perturbation by dye addition: |  |  |  | 0.0027±0.0022
| Stability constants: |  |  |  | Dickson, A.G., Sabine, C.L., and Christian, J.R. (Eds.) 2007. Guide to best practices for ocean CO2 measurements. PICES Special Publication 3, 191 pp.
Spectrophotometer model: | [5 empty] | Agilent 8543
| Accuracy: |  |  |  |  | Peak wavelengths confirmed to within 1 nm by spectrophotometric scans of a sealed HoO standard
| Stability: |  |  |  |  | Self tests were run and passed
| Absorbance change at 730 nm w/dye addition: |  |  |  |  | 0.004±0.014
Notes: | 34 | Unstable analysis; inconsistent w/TIC & Alk
| 37 | Unstable background; inconsistent w/TIC & Alk
| 44 | Inconsistent w/TIC & Alk
| 204 | Unstable analysis, large difference between replicates
| 211 | C parameters inconsistent
| 316 | Unstable analysis; inconsistent w/TIC & Alk
| 318 | C parameters inconsistent; Unstable background
| 376 | Large difference between replicates
| 381 | Unstable analysis; inconsistent w/TIC & Alk
| 384 | Unstable analysis; inconsistent w/TIC & Alk
| 558 | Large difference between replicates
Replicate Analyses (same niskin)
Sample ID | A | B | Average | difference | abs diff
30 | 7.3077 | 7.3042 | 7.306 | 0.0034286 | 0.0034286
195 | 7.4112 | 7.4097 | 7.4105 | 0.0015584 | 0.0015584
204 | 7.5804 | 7.5922 | 7.5863 | -0.011845 |  | Unstable analyses
209 | 7.7603 | 7.7634 | 7.7618 | 0.0031608 | 0.0031608
309 | 7.5215 | 7.5278 | 7.5247 | 0.0063328 | 0.0063328
322 | 7.4423 | 7.4399 | 7.4411 | -0.002399 | 0.002399
327 | 7.7312 | 7.7331 | 7.7322 | 0.0019564 | 0.0019564
363 | 7.5157 | 7.5256 | 7.5157 | -0.0099201 |  | Unstable analyses
376 | 7.5328 | 7.5423 | 7.5376 | -0.0094659 | 0.0094659
381 | 7.7504 | 7.7511 | 7.7507 | 0.00068412 | 0.00068412
429 | 7.5413 | 7.5313 | 7.5313 | 0.0099802 |  | Unstable background
441 | 7.3884 | 7.3891 | 7.3888 | -0.0006706 | 0.0006706
558 | 7.3888 | 7.3965 | 7.3926 | -0.0077062 | 0.0077062
559 | 7.3914 | 7.3921 | 7.3917 | -0.0007156 | 0.0007156
560 | 7.3924 | 7.3915 | 7.392 | 0.00088674 | 0.00088674
561 | 7.3891 | 7.3899 | 7.3895 | -0.00081713 | 0.00081713
562 | 7.3905 | 7.3951 | 7.3928 | -0.0045805 | 0.0045805
average variability: |  |  |  | -0.0011842 | 0.0031688
standard deviation of variability: |  |  |  | 0.0059223 | 0.0028574
Replicate Samples (different niskins)
|  | Average | stdev
558 | 7.3926
559 | 7.3917
560 | 7.392
561 | 7.3895
562 | 7.3928
|  | 7.3911 | 0.0013285
Precision estimate: |  | 0.0032
Version control
2009-08-04 | First archived
2010-05-18 | Temperature and nutrient data added; raw data removed