CHEMICAL DATA BASE METHODS

(Instituto de Investigacións Marinas. CSIC. Spain)

Aida F. Ríos

aida@iim.csic.es

A Metrohm E-654 pH-meter equipped with a Ross (Orion 81-04) combined glass electrode was used to determine pH on the NBS scale. The temperature was measured using a platinum resistance thermometer and finally pH was referred to a standard temperature of 15°C (pH15) according to Pérez and Fraga (1987a). The method has a shipboard precision of ±0.002 pH15 (Ríos and Rosón, 1996) and an accuracy of ±0.004 pH15 using samples of Certified Reference Material (CRMs) provided by Dr. Dickson from the Scripps Institution of Oceanography (Ríos and Pérez, 1999; Ríos and Rellán, 1998).

Alkalinity was determined by automatic potentiometric titration with HCl at a final pH of 4.44 (Pérez and Fraga, 1987b). The electrodes were standardised using an NBS buffer of pH 7.413 and checked using an NBS buffer of 4.008. This method has a precision of 0.1% (Pérez and Fraga, 1987b), and an accuracy of 1.4µmol.kg-1 (Ríos and Pérez, 1999; Ríos and Rellán, 1998).

## Carmen G. Castro

ccarmen@iim,csic.es

Nutrient samples were filtered through 0.45 µm Millipore filter prior to analysis and were analysed within 12 h after collection; and were stored in the refrigerator prior to analysis and in the dark. Nutrient concentrations were determined by segmented flow analysis with Technicon AAII systems, following Hansen and Grasshoff (1983) with some improvements (Mouriño and Fraga, 1985; Álvarez-Salgado et al., 1992). The analytical error was 0.05 µmolkg-1 for nitrate, 0.05 µmolkg-1 for silicic acid and 0.01 µmolkg-1 for phosphate. Dissolved oxygen was determined by Winkler potentiometric titration. The estimated analytical error was 1 µmolkg-1. Oxygen saturation was calculated following Benson and Krause equation (UNESCO, 1986). Chlorophyll a was measured using 90% acetone extraction in a 10,000 R Turner fluorometer (Yentsch and Menzel, 1963). The precision was  0.05 mgm-3. Particulate organic matter (filtration volume 1 l) was collected on Whatman GF/F filters and analyses were performed in a PE 2400 elemental analyser, with a precision of 0.04 µmolkg-1 for nitrogen and 0.1 µmolkg-1 for carbon.

M.D. Doval

marylo@iim.csic.es, mdoval@cccmm.cesga.es

Seawater samples for DOC analysis were collected with 100 ml polyethylene syringes with teflon plunger tips and filtered by hand through Whatman Puradisc GF/F disposable filter devices (0.7µm pore size) on polypropylene housing. The filtrate was drawn eventually into 50 ml polyethylene containers. The filtering system and the containers used for DOC had been previously soaked on 0.1 N HCl, and rinsed with Milli-Q water. In addition, the containers were rinsed three times with 50 ml of sample. Samples were immediately stored at -70ºC until analysis in the base laboratory, eight months later. This storage technique has demonstrated no artefactual results on the micromolar scale (Hansell and Carlson, 1998b).

DOC determination was performed by high temperature catalytic oxidation (HTCO) with a commercial Shimadzu TOC-5000. The combustion quartz tube was filled with a 0.5% Pt on Al2O3 catalyst. Three to 5 replicate injections of 200 µl were performed per sample. The concentration of DOC was determined by subtracting the average peak area from the instrument blank area and dividing by the slope of the standard curve. The instrument blank is the system blank plus the filtration blank. The system blank was determined by subtracting the DOC in UV-Milli-Q to the total blank. Measurements made with the high sensitivity catalyst (Pt on silica wool) produced values <2 µmol C l-1 for fresh UV-Milli-Q water. The filtration blank (determined by filtering UV-Milli-Q water through the filtration system) was <2 µmol C l-1. Before sample analyses, the catalyst was washed by injecting UV-Milli-Q, for at least 12 h, until the system blank was low and stable. The system blank was <8 µmol C l-1. The device was standardized with Potassium Hydrogen Phthalate (KHP). The coefficient of variation (C.V.) of the peak area for the 3-5 replicate analyses of each sample was ~1%. The accuracy of our HTCO system has been tested within the international intercalibration exercise conducted by J. Sharp (Univ. of Delaware), with very satisfactory results (within 10%; J. Sharp, pers. com.).