HLY-04-03 Service Group Bottle Data Documentation

18 July to 26 August 2004

Nome, Alaska to Nome, Alaska

Dr. Louis Codispoti, (on board PI)

Horn Point Laboratory

PO Box 775

2020 Horn Pt. Rd.

Cambridge, MD 21613

410.221.8479 phone

codispot@hpl.umces.edu

Dr. James Swift (PI)

Scripps Institution of Oceanography

Oceanographic Data Facility

9500 Gilman Rd. MC 0214

La Jolla, CA 92093-0214

858.534.3387 phone

jswift@ucsd.edu

On board team: Kristin Sanborn, Rebekah Duncan, Dan Schuller, Erik Quiroz,

Robert Palomares

Other team members: Berit Rabe (ADCP), Dean Stockwell (Chlorophyll)

Funded by National Science Foundation OPP0125399

Data Set Overview

184 CTD casts on 60 stations were attempted. 4 of these were aborted, with no CTD data and no water samples.  These casts were Station 022 Cast 01, Station 022 Cast 03, Station 022 Cast 04, and Station 049 Cast 01.

# Instrumentation

CTD casts were performed with a rosette system consisting of a 12-place rosette frame with 30 liter bottles and a 12-place SBE-32 Carousel pylon.  Underwater electronic components consisted of:

- Sea-Bird Electronics, Inc. (SBE) 911plus CTD,

- WETLabs C-Star transmissometer with a 25cm path length and 660nm wavelength,

- Biospherical Instruments, Inc. Photosynthetically Active Radiation (PAR) sensor,

- Chelsea MkIII Aquatracka fluorometer, and

- Simrad, 5 volt – 2-500 meter altimeter.

Additionally, a Dr. Haardt fluorometer designed to detect colored organic matter (CDOM) and a Secchi disk were mounted on the CTD package. The CTD, transmissometer, and the two fluorometers were mounted horizontally along the bottom of the rosette frame. The PAR sensor was located at the top of the rosette. The surface PAR sensor was located on a newly installed I-beam antenna support atop the helicopter shack.  All sensors except the Secchi disk were interfaced with the CTD system. This instrument package provided pressure, dual temperature and dual conductivity channels as well as light transmissivity and fluorometric signals at a sample rate of 24 scans per second.

The bottles on the rosette were General Oceanic 30 liter bottles. The bottles were equipped with internal nylon coated springs and silicone o-rings which are used to minimize toxicity to the sample. Bottle numbering is 1 to 12 with 1 tripped first usually at the deepest sampling level and 12 tripped last at the shallowest sampling level. The rosette system was suspended from a standard UNOLS 3 conductor 0.322” electromechanical cable.

The CTD used was serial number 09P24152-0638 and the sensor’s model and serial numbers are listed in Table 1. At Station 022, cast 1, the CTD failed when three connectors were found flooded. The CTD was replaced with 09P24152-0639. After cast 8 at Station 022, it was also determined that the Haardt fluorometer was damaged and it was removed from service. This data from Station 022 has been deleted.

TABLE 1. Instrument/Sensor Serial Numbers

| Primary Temperature | Primary Conductivity | Secondary Temperature | Secondary Conductivity | Pressure | Transmissometer |
| --- | --- | --- | --- | --- | --- |
| SBE 3plus | SBE 4C | SBE 3plus | SBE 4C | 401K-105 | C-Star |
| 03-2796 | 04-2545 | 03-2824 | 04-2568 | 83009 | CST-390DR |

| Oxygen | Fluorometer | PAR | Surface PAR | Altimeter |
| --- | --- | --- | --- | --- |
| SBE 43 | Aqua 3 | QSP-2300 | QSR-240 | 807 |
| 0459 | 088233 | 4643 | 6367 | 9711090 |

## Equipment Positions

TABLE 2.  Instrument mounting heights in reference to the bottom of the rosette frame.

| Sensor | Height above base of rosette |  | Sensor | Height above base of rosette |
| --- | --- | --- | --- | --- |
| Altimeter | 2 cm |  | Pressure | 19cm |
| Transmissometer | 8 cm |  | T (pri) | 10 cm |
| Fluorometer (Chelsea) | 10 cm |  |  |  |
| Fluorometer (Haardt) | 8 cm |  | Par | 215cm Sta. < 2000 m |

The distance of the mid-points of the 30 L Niskin bottles from the bottom-mounted sensors was ~1.19m. The 30 Liter Niskin bottles are ~1.0 m long. The secchi disk was mounted 2.2m above the bottom of the rosette frame.

## Problems and/or Procedural changes

At times small areas of the nylon coating on the Niskin bottle springs broke down so there may be some minor iron contamination in some samples. HLY0403 rosette operations were continually beset by problems with bottle leaks caused by Niskin bottle end o-rings falling out of position.  Typically, each cast had one such occurrence.  Data from leaking bottles were carefully assessed by comparing CTD oxygen and salinity data against bottle salinities and temperatures and it was found that spigot leaks normally did not compromise sample quality. Samples were usually not taken from bottles with severe leaks caused by poorly sealed bottom end-caps, and in other cases data from leaking bottles were noted as being questionable.  Thus, the main effect of the leaking bottles was to increase the sample spacing of good data on some casts. Some bottles were more prone than others to o-ring problems, but no clear-cut patterns emerged except that there seemed to be more problems on deeper casts which would implicate a pressure effect or prolonged exposure to low temperatures.  Since both the bottles and o-rings were purchased just before the two 2004 Healy cruises, we suspect a manufacturing defect or change in either the bottles or o-rings. Some of the problems were gross, i.e. the o-ring would be visible out the side of the end cap, but others were more subtle. Every time an o-ring problem was suspected, the o-ring was carefully inspected, and replaced if necessary.  Also, at several points during the cruise all o-rings were inspected.  During the spring expedition, HLY-04-02, the contents of various packages of spare o-rings were measured to locate 'large' or 'small' o-rings (within the manufacturer's tolerance), and a remedial 'large' set was installed.  Another time Coast Guard personnel replaced all the o-rings from their own supply.  Yet all these remedial attempts were to no particular avail. A different batch of o-rings were procured prior to this cruise (HLY-04-03) but they  made minimal difference vis a vis  bottle leaks.

# CTD Data

## CTD Laboratory Calibration Procedures

Pre-cruise laboratory calibrations of CTD pressure, temperature and conductivity sensors were used to generate coefficients for the calculation of these parameters from their respective sensor frequencies. The temperature, conductivity and oxygen sensors were recalibrated prior to the HLY-04-3 expedition. The temperature, conductivity and oxygen calibrations were performed at Sea-Bird Electronics, Inc. in Bellevue, Washington.  Calibration of the pressure sensor was performed prior to the HLY0402 spring cruise by Scripps Institution of Oceanography, Shipboard Technical Support/Oceanographic Data Facility (SIO/STS/ODF) personnel. The Sea-Bird laboratory temperature calibrations were referenced to the International Temperature Scale of 1990 (ITS-90).

## CTD Data Acquisition

The CTD 911plus was operated generally as suggested in the Sea-Bird CTD Operating and Repair Manual, which contains a description of the system, its operation and functions (Sea-Bird Electronics, Inc., 2002). One difference from Sea-Bird’s operation is that data acquisition was started on deck. This procedure allows a check of the pressure offset and an unblocked reading of the transmissometer. The Seasoft acquisition program as described in the CTD Data Acquisition Software Manual (Sea-Bird Electronics, Inc., 2001) provided a real-time graphical display of selected parameters adequate to monitor CTD performance and information for the selection of bottle-tripping depths. Raw data from the CTD were archived on the PC’s hard disk at the full 24 Hz sampling rate.

A CTD Station Sheet form was filled in for each deployment, providing a record of times, positions, bottom depth, bottle sampling depths, and every attempt to trip a bottle, as well as any pertinent comments. When the equipment and personnel were ready, data acquisition was started.  The CTD operator pressed a control key (flag), which appends a summary line into the files created for “inventory” files.  This file contains a summary of the time, ship’s position, and current scan number each time the control key is pressed.  They are used as a reference to mark important events during the cast, such as on deck pressure, when the lowering was initiated, when the package was at the bottom, when bottles were tripped and the on-deck pressure with ending position.  After the initial flag, the rosette/CTD system was lowered into the water and held at 5 meters wire out for 3-5 minutes to permit activation of the CTD pumps and equilibration of the sensors.  Then, the operator had the CTD raised to the surface, again created a flag, and simultaneously directed the winch operator to begin lowering.  The operator created a flag at the deepest point of the cast. Bottom depths were calculated by combining the distance above bottom, reported by the altimeter, and the maximum depth of the CTD package when bottom altimeter readings were available.  If there was no altimeter reading, then the bottom depth is reported from the ship’s Bathy 2000 or Knudsen model 320B/R depth recorder.  These data, corrected for the draft of the transducer, were logged in uncorrected meters (assuming a sound velocity of 1500 m/sec). If the altimeter and depth recorder data were unavailable, the final resort was to use depth data from the SeaBeam system (corrected sound velocities).

The wire out corresponding to each bottle trip was written on the station log and the trips were electronically flagged in the data file.  The performance of all sensors was monitored during the cast.  After the rosette recovery, the operator created a final flag denoting the end of the cast. Any faulty equipment or exceptionally noisy data were noted on the log sheet.

## CTD Data Processing

### Pressure

CTD values determined on deck before and after each cast were compared to determine a pressure offset correction. The comparison suggested that no pressure offset was necessary.

### Temperature

The temperature sensors were calibrated in July 2004.  The dual temperature sensors were monitored during the expedition and exhibited good agreement.  It appears that no additional corrections need to be applied. A post-cruise calibration will be performed.

### Conductivity

The conductivity sensors were calibrated in July 2004. Corrected CTD pressure and temperature values were used with bottle salinities to back-calculate bottle conductivities. Comparison of these bottle values with the CTD primary conductivity values indicated no additional offset needed to be applied to the data.

### Transmissometer

A WETLabs calibrated transmissometer was utilized throughout the cruise.  An on deck calibration check was performed and even though there was little degradation from the last calibration, the new coefficients were applied to the data set.

### Oxygen, Fluorometer, and PAR

The CTD oxygen data are only intended for qualitative use.  The oxygen sensor was calibrated in July 2004. Similarly, the fluorometric and PAR data are not calibrated.

### Data Processing

Sea-Bird Seasoft CTD processing software was employed. The processing programs are outlined below.  A more complete description may be found in the Sea-Bird Software Manual which is available from the Sea-Bird website (www.seabird.com).

The sequence of programs that were run in processing CTD data from this cruise are as follows:

- DATCNV - Converts data from raw frequencies and voltages to corrected engineering units

- WILDEDIT - Eliminates large spikes

- CELLTM - Applies conductivity cell thermal mass correction

- FILTER – A low pass filter to smooth pressure for LOOPEDIT

- LOOPEDIT - Marks scans where velocity is less than selected value to avoid pressure reversals from ship roll, or during bottle flushing.

- DERIVE - Computes calculated parameters

- BINAVG - Average data into desired pressure bins

The quality control steps included:

- Sensor verification After the CTD was set up and sensor serial numbers and sensor location entered into the computer, another check was made to verify that there were no tabulation errors.

- Seasoft Configuration File was reviewed to verify that individual sensors were represented correctly, with the correct coefficients.

- Temperature was verified by comparing primary and secondary sensor data.

- Conductivity was checked by comparison of the two sensors with each other and with bottle salinity samples.

- Position Check A chart of the ship’s track was produced and reviewed for any serious problems.  The positions were acquired from the ship’s GPS (P-code) navigation system.

- Visual Check Plots of each usable cast were produced and reviewed for any noise and spikes that may have been missed by the processing programs.

- The density profile was checked for inversions that might have been produced by sensor noise or response mismatches.

### CTD Data Footnoting

WHP water bottle quality flags were assigned as defined in the WOCE Operations Manual (Joyce and Corry, 1994). These flags and interpretation are tabulated in the CTD and Bottle Data Distribution, Quality Flags section of this document.

### Data Comments

Fine structure including minor density inversions that may appear in the upper ~ 10 m of the profiles is most likely caused by ship discharges/turbulence. To minimize the ship effect, engine cooling water discharges were restricted to the port side of the Healy.  A “yo yo” procedure was adopted to induce bottle flushing whenever waves and ship motion were weak.  This procedure was employed for all bottle trips under quiescent conditions except for productivity casts, and for surface and bottom bottles that were usually kept at depth for ~2 minutes before tripping whereas the yo yo procedure generally involved a ~30 second wait, a yo yo and another ~30 second wait. Regardless of the procedure employed, the CTD operators were instructed to wait for at least 1 minute (typically > 1.5 minutes) before tripping the bottle.

All salinity, nutrient and dissolved oxygen data collected by the “service” team have gone through several stages of editing and are not likely to change significantly.  The chlorophyll observations reported here may may undergo further editing.  Please note that changes in the refractive index of samples with salinity, baseline issues, temperature effects, etc. can cause variability in the ammonium and urea methods that is not reflective of changes in ambient concentrations. Differences of < ~ 0.15 µM should be treated with caution. Please contact L. A. Codispoti, if you wish to interpret small differences in ammonium and urea data (Codispoti@hpl.umces.edu) The other nutrient data meet WOCE/JGOFS standards.

# Bottle Data

There were five generic types of casts performed with differing sampling protocols.  Generally speaking, the samplings associated with these casts were as follows, but there is some cast-to-cast variation.

- Hydrographic

  - Oxygen,

  - Total CO2,

  - Total Alkalinity,

  - Nutrients

  - Chlorophyll/Phaeophytin

  - Phytoplankton

  - Salinity

  - O18/O16

  - Benthic

  - Dissolved Organic Matter/Particulate Organic Matter

  - Thorium-234

  - C13/N15

- Productivity/Zooplankton

  - Oxygen

  - Oxygen Respiration

  - Productivity

  - Nutrients

  - Chlorophyll

  - HPLC

  - Bacteria

  - Micro Zooplankton

  - Particulate Organic Matter

  - Dissolved Organic Matter/Lignin

  - Bio-Optics

  - Taxonomy

  - C13/N15

- Bio-Markers

  - Nutrients

  - Particulate Organic Matter

  - Dissolved Organic Matter/Lignin

  - C13/N15

- Radium

  - Nutrients

  - Radium

  - C13/N15

- Zooplankton

  - Nutrients

  - Micro Zooplankton

  - C13/N15

The correspondence between individual sample containers and the rosette bottle from which the sample was drawn was recorded on the sample log for the cast. This log also included any comments or anomalous conditions noted concerning  the rosette and bottles.

Normal sampling practice included opening the drain valve before the air vent on the bottle, to check for air leaks. This observation and other diagnostic comments (e.g., "lanyard caught in lid", "valve left open") that might later prove useful in determining sample integrity were routinely noted on the sample log.

## Bottle Data Processing

After the samples were drawn and analyzed, the next stage of processing involved merging the different data streams into a common file. The rosette cast and bottle numbers were the primary identification for all ODF-analyzed samples taken from the bottle, and were used to merge the analytical results with the CTD data associated with that bottle.

Diagnostic comments from the sample log, and notes from analysts and/or bottle data processors were entered into a computer file associated with each station (the "quality" file) as part of the quality control procedure. Sample data from bottles suspected of leaking were checked to see if the properties were consistent with the profile for the cast, with adjacent stations, and, where applicable, with the CTD data. Direct inspection of the tabular data, property-property plots and vertical sections were all employed to check the data. Revisions were made whenever there was an objective reason to delete, annotate or re-calculate a datum. WHP water sample codes were selected to indicate the reliability of the individual parameters affected by the comments. WHP bottle codes were assigned where evidence showed the entire bottle was affected, as in the case of a leak, or a bottle trip at other than the intended depth.

Specific data processing and techniques and additional quality control are included with the parameter write-up.

### Pressure and Temperatures

All pressures and temperatures for the bottle data tabulation were obtained by averaging CTD data for a brief interval at the time the bottle was closed and then applying the appropriate calibration data.

The temperatures are reported using the International Temperature Scale of 1990.

### Salinity

337 salinity samples were analyzed in 13 analyses runs.

#### Sampling and Data Processing

Salinity samples were drawn into 200 ml high alumina borosilicate bottles, which were rinsed three times with sample prior to filling. The bottles were sealed with custom-made plastic insert thimbles and Nalgene screw caps. This container provides very low container dissolution and sample evaporation.

#### Equipment and Techniques

A Guildline Autosal 8400B #65-715, standardized with IAPSO Standard Seawater (SSW) batch P-144, was used to measure the salinities. Prior to the analyses, the samples were stored to permit equilibration to laboratory temperature, usually 8-20 hours.  The salinometer was outfitted with an Ocean Scientific International interface for computer-aided measurement. The salinometer was standardized with a fresh vial of standard seawater (SSW) at the beginning of each analysis run.  Instrument drift was determined by running a SSW vial after the last sample was run through the autosal. The salinometer cell was flushed until two successive readings met software criteria for consistency; these were then averaged for a final result. The estimated accuracy of bottle salinities run at sea is usually better than 0.002 PSU relative to the particular standard seawater batch used.

#### Laboratory Temperature

The temperature stability in the salinometer laboratory was moderate; variation was no more than 3ºC during a run of samples.  The laboratory temperature was generally 2-3ºC lower than the Autosal bath temperature.

### Oxygen

974 samples were analyzed for oxygen.

#### Sampling and Data Processing

Samples were collected for dissolved oxygen analyses as the first sample after the rosette was brought on board. Using a Tygon drawing tube, nominal 125ml volume-calibrated iodine flasks were rinsed then filled and allowed to overflow for approximately 3 flask volumes. The sample draw temperature was measured with a small platinum resistance thermometer embedded in the drawing tube. Reagents were added to fix the oxygen before stoppering. The flasks were shaken twice to assure thorough dispersion of the precipitate, once immediately after drawing, and then again after about 20 minutes.  The samples were usually analyzed within several hours of collection.

Thiosulfate normalities were calculated from each standardization and corrected to 20ºC. Periodically, the 20ºC normalities and the blanks were plotted versus time and were reviewed for possible problems. New thiosulfate normalities were recalculated as a linear function of time, if warranted. The oxygen data were recalculated using the smoothed normality and an averaged reagent blank. Oxygens were converted from milliliters per liter to micromoles per kilogram using the sampling draw temperature and the sample’s salinity.

#### Equipment and Techniques

Dissolved oxygen analyses were performed with an ODF-designed automated oxygen titrator using photometric end-point detection based on the absorption of 365nm wavelength ultra-violet light. The titration of the samples and the data logging were controlled by PC software. Thiosulfate was dispensed by a Dosimat 665 buret driver fitted with a 1.0 ml buret. The ODF method used a whole-bottle modified-Winkler titration following the technique of Carpenter (1965) with modifications by Culberson (1991), but with higher concentrations of potassium iodate standard (approximately 0.012N) and thiosulfate solution (55 g/l). Standard KIO3 solutions prepared ashore were run at the beginning of each run. Reagent and distilled water blanks were determined, to account for presence of oxidizing or reducing materials.

#### Volumetric Calibration

Oxygen flask volumes were determined gravimetrically with degassed deionized water to determine flask volumes at ODF’s chemistry laboratory. This was done once before using flasks for the first time and periodically thereafter when a suspect bottle volume was detected. The volumetric flasks used in preparing standards were volume-calibrated by the same method, as was the 10 ml Dosimat buret used to dispense standard iodate solution.

#### Standards

Potassium iodate was obtained from Johnson Matthey Chemical Co. and was reported by the supplier to be >99.4% pure.

### Nutrients

1567 samples were analyzed for nutrients in 70 analyses runs. 68 of these samples were checks of the system and 177 were for 5 of the ancillary shipboard programs.

#### Sampling and Data Processing

Nutrient samples were drawn into 45 ml polypropylene, screw-capped “oak-ridge type” centrifuge tubes. The tubes were rinsed with 10% HCl and then with sample three times before filling. Standardizations were performed at the beginning and end of each group of analyses (4-38 samples) with an intermediate concentration mixed nutrient standard, which was prepared prior to each run from a secondary standard in a low-nutrient seawater matrix. The secondary standards were prepared aboard ship by dilution from primary standard solutions.  Dry standards were pre-weighed at the laboratory at ODF, and transported to the vessel for dilution to the primary standard. Sets of 7 different standard concentrations covering the range of sample concentrations were analyzed periodically to determine the deviation from linearity, if any, as a function of absorbance for each nutrient analysis.  A correction for non-linearity was applied to the final nutrient concentrations when necessary.

After each group of samples was analyzed, the raw data file was processed to produce another file of response factors, baseline values, and absorbances.  Suspect values were checked for accuracy against values taken from strip chart recordings. Sargasso Sea Water and deep seawater were run occasionally to check ammonium and urea baselines and assess system performance in all six nutrient channels.

Nutrients, when reported in micromoles per kilogram, were converted from micromoles per liter by dividing by sample density calculated at 1 atm pressure (0 db), in situ salinity, and the sample temperature measured at the time of analysis.

Also reported is N**, a parameter calculated from nitrate, nitrite, ammonium and phosphate concentrations.  This parameter is defined as N** = ((N-16P + 2.98)µM) 0.87, where P = the phosphate concentration in µM, and N = (nitrate+nitrite+ammonium in µM).  This parameter is quite similar to the original N* parameter defined by Gruber and Sarmiento (1997) except that we include ammonium concentrations because of the high ammonium concentrations that can occur in the SBI region.  The underlying premise of both N* and N** is that the N/P atomic regeneration ratio in seawater is normally close to the 16/1 N/P Redfield ratio.  The assumption is that deviations from this ratio in N/P ratios in a water mass arise primarily from nitrogen fixation which produces organic matter with N/P ratios in excess of 16/1, or denitrification which consumes nitrate and other forms of fixed nitrogen and converts these forms into elemental dinitrogen gas.  Values less than 2.98 suggest that a water mass has experienced net denitrification and higher values suggest net nitrogen fixation.  The factors 2.98 and 0.87 are explained by Gruber and Sarmiento (1997), and there is some debate about whether they should be included, but we do so in order to facilitate comparison with the distributions presented by Gruber and Sarmiento (1997).

#### Equipment and Techniques

Nutrient analyses (nitrate+nitrite, nitrite, phosphate, silicate, ammonium, and urea) were performed on an ODF-modified 6-channel Technicon AutoAnalyzer II, generally within a few hours after sample collection.  The samples were kept in the dark by covering with tin foil or refrigerated at 4C, if necessary, but brought to within 5C of lab temperature before analysis.  The analog outputs from each of the six channels were digitized and logged automatically by computer (PC) at 2-second intervals.

A modification of the Armstrong et al. (Armstrong 1967) procedure was used for the analysis of nitrate and nitrite. For the nitrate plus nitrite analysis, the seawater sample was passed through a cadmium reduction column where nitrate was quantitatively reduced to nitrite. The stream was then passed through a 15mm flowcell and the absorbance measured at 540nm.  The same technique was employed for nitrite analysis, except the cadmium column was bypassed, and a 50mm flowcell was used for measurement.  Periodic checks of the column efficiency were made by running alternate equal concentrations of NO2 and NO3 through the NO3 channel to ensure that column efficiencies were high (> 95%). Nitrite concentrations were subtracted from the nitrate+nitrite values to obtain nitrate concentrations.

Phosphate was analyzed using a modification of the Bernhardt and Wilhelms [Bernhardt 1967.] technique. The reaction product was heated to ~55ºC to enhance color development, then passed through a 50mm flowcell and the absorbance measured at 820m.

Silicate was analyzed using the technique of Armstrong et al., (Armstrong, 1967). The sample was passed through a 15mm flowcell and the absorbance measured at 660nm.

Ammonium was determined by the Berthelot reaction (Patton and Crouch 1977) in which sodium hypochlorite and phenol react with ammonium ion to produce indophenol blue, a blue compound.   The solution was heated to 55°C and passed through a 50mm flowcell at 640nm.

Urea was analyzed via a modification of the method of Rahmatullah and Boyde (1980), which is based on the classic diacetyl monoxime method.  A solution of diacetyl monoxime, thiosemicarbizide and acetone is followed by the addition of ferric chloride, which acts as a catalyst.  The resultant solution was heated to 90°C and passed through a 50mm flowcell. The absorbance was measured at 520nm.

#### Nutrient Standards

Na2SiF6, the silicate primary standard, was obtained from Johnson Matthey Company and Fisher Scientific and was reported by the suppliers to be >98% pure. Primary standards for nitrate (KNO3), nitrite (NaNO2), and phosphate (KH2PO4) were obtained from Johnson Matthey Chemical Company, and the supplier reported purities of 99.999%, 97%, and 99.999%, respectively. Ammonia, (NH4(SO4)2), and Urea primary standards were obtained from Fisher Scientific and reported to be >99% pure.

### Bottle Data Footnoting

WHP water bottle quality flags were assigned as defined in the WOCE Operations Manual [Joyce]. These flags and interpretation are tabulated in the Data Distribution, Bottle Data, Quality Flags section of this document.

# Data Distribution

The CTD and bottle data can be obtained through NCAR’s Earth Observing Laboratory website, www.eol.ucar.edu/projects/sbi. The data are reported using the WHP-Exchange (WOCE Hydrographic Program) format and the quality coding follows those outlined by the WOCE program (Joyce, 1994). In addition, the format can be obtained through the WOCE Hydrographic Program website, WHPO.ucsd.edu. The descriptions in this document have been edited from the reference to annotate the format specific to this data distribution. ASCII files for each station were created with comments recorded on the CTD Station Logs during data acquisition. These ASCII files include data processing comments noting any problems, their resolution, and footnoting that may have occurred.  A separate ASCII file was also created with the comments from the Sample Log Sheets that include problems with the Niskin bottles that could compromise the samples. Comments arising from inspection and checking of the data are also included in the ASCII file. These comment files are also in the EOL/JOSS database.  Raw (unprocessed) CTD data files are archived in the EOL/JOSS database.  The file hly0403_ctd_raw.zip contains ssscc.cfg, ssscc.con, ssscc.dat and ssscc.hdr (where sss = station number and cc = cast number) files as acquired by the SeaBird SeaSave acquisition program, sbscan.sum file and calibration information for all sensors.  The *.cfg file is datcnv.cfg with the beginning scan number and *.con files may include a correction based on the bottle salinity samples.  The sbscan.sum file is a list of stations and beginning scan number.  Configuration files for the various SeaBird CTD processing programs are also included where applicable.

## General rules for WHP-exchange:

- Each line must end with a carriage return or end-of-line.

- With the exception of the file type line, lines starting with a "#" character, or including and following a line which reads "END_DATA", each line in the file must have exactly the same number of commas as do all other lines in that file.

- The name of a quality flag always begins with the name of the parameter with which it is associated, followed by an underscore character, followed by "FLAG", followed by an underscore, and then followed by an alphanumeric character, W.

- The "missing value" for a data value is always defined as -999, but written in the decimal place format of the parameter in question. For example, a missing salinity would be written -999.0000 or a missing phosphate -999.00.

- The first four characters of the EXPOCODE are the U.S. National Oceanographic Data Center (NODC) country-ship code, then followed by up to an 8 characters expedition name of cruise number, i.e. 32H1HLY0403.

## CTD Data

### CTD data is located in file 32H1hly0403_ct1.zip.  This file contains ssscc_ct1.csv files for each station and cast where sss=3 digit station identifier and cc=2 digit cast identifier.

### Description of ssscc_ct1.csv file layout.

| 1st line | File type, here CTD, followed by a comma and a DATE_TIME stamp YYYYMMDDdivINSwho YYYY 4 digit year MM 2 digit month DD 2 digit day div division of Institution INS Institution name who initials of responsible person |
| --- | --- |
| # lines | A file may include 0-N optional lines at the start of a data file, each beginning with a "#" character and each ending with carriage return or end-of-line. Information relevant to file change/update history may be included here, for example. |
| 2nd line | NUMBER_HEADERS = n (n = 10 in this table and the example_ct1.csv file.) |
| 3rd line | EXPOCODE = [expocode] The expedition code, assigned by the user. |
| 4th line | SECT_ID = [section] The SBI station specification. Optional. |
| 5th line | STNNBR = [station] The originator's station number |
| 6th line | CASTNO = [cast] The originator's cast number |
| 7th line | DATE = [date] Cast date in YYYYMMDD integer format. |
| 8th line | TIME = [time] Cast time that CTD was at the deepest sampling point. |
| 9th line | LATITUDE = [latitude] Latitude as SDD.dddd where "S" is sign (blank or missing is positive), DD are degrees, and dddd are decimal degrees. Sign is positive in northern hemisphere, negative in southern hemisphere |
| 10th line | LONGITUDE = [longitude] Longitude as SDDD.dddd where "S" is sign (blank or missing is positive), DDD are degrees, and dddd are decimal degrees. Sign is positive for "east" longitude, negative for "west" longitude |
| 11th line | DEPTH = [bottom] Reported depth to bottom. Preferred units are "meters" and should be specified in Line 2. In general, corrected depths are preferred to uncorrected depths. Documentation accompanying data includes notes on methodology of correction. Optional. |
| next line | Parameter headings. |
| next line | Units. |
| data lines | A single _ct1.csv CTD data file will normally contain data lines for one CTD cast. |
| END_DATA | The line after the last data line must read END_DATA, and be followed by a carriage return or end of line. |
| other lines | Users may include any information they wish in 0-N optional lines at the end of a data file, after the END_DATA line. |

### Parameter names, units, format, and comments

| Parameter | Units | Format | Comments |
| --- | --- | --- | --- |
| CTDPRS | DB | F7.1 | CTD pressure, decibars |
| CTDPRS_FLAG_W |  | I1 | CTDPRS quality flag |
| CTDTMP | ITS-90 | F8.3 | CTD temperature, degrees C (ITS-90) |
| CTDTMP_FLAG_W |  | I1 | CTDTMP quality flag |
| CTDSAL |  | F8.3 | CTD salinity |
| CTDSAL_FLAG_W |  | I1 | CTDSAL quality flag |
| CTDOXY | UMOL/KG | F7.1 | CTD oxygen, micromoles/kilogram |
| CTDOXY_FLAG_W |  | I1 | CTDOXY quality flag |
| STHETA |  | F8.3 | Sigma Theta |
| STHETA_FLAG_W |  | I1 | Sigma Theta quality flag |
| XMISS | %TRANS | F7.1 | Transmissivity, percent transmittance |
| XMISS_FLAG_W |  | I1 | XMISS quality flag |
| FLUOR | VOLTS | F8.3 | Fluorometer, voltage |
| FLUOR_FLAG_W |  | I1 | Fluorometer quality flag |
| PAR | VOLTS | F8.3 | PAR, voltage |
| PAR_FLAG_W |  | I1 | PAR quality flag |
| SPAR | VOLTS | F8.3 | Surface PAR, voltage |
| SPAR_FLAG_W |  | I1 | Surface PAR quality flag |
| FLCDOM | VOLTS | F8.3 | CDOM Fluorometer, voltage |
| FLCDOM_FLAG_W |  | I1 | CDOM Fluorometer quality flag |
| DEPTH | METERS | F8.0 | Depth |

### Quality Flags

CTD data quality flags were assigned to the CTDTMP (CTD temperature), CTDSAL (CTD salinity) and XMISS (Transmissivity) parameters as follows:

- Acceptable measurement.

- Questionable measurement. The data did not fit the station profile or adjacent station comparisons (or possibly bottle data comparisons). The data could be acceptable, but are open to interpretation.

- Bad measurement. The CTD data were determined to be unusable.

- Not reported. The CTD data could not be reported, typically when CTD salinity is flagged 3 or 4.

- Not sampled. No operational sensor was present on this cast

WHP CTD data quality flags were assigned to the CTDOXY (CTD O2), FLUORO (Fluorometer), PAR (PAR), SPAR (Surface PAR), and HAARDT (Haardt Fluorometer CDOM) parameter as follows:

- Not calibrated. Data are uncalibrated.

- Not sampled. No operational sensor was present on this cast. Either the sensor cover was left on or the depth rating necessitated removal.

## Bottle Data

### Description of 32H1HLY0403_hy1.csv file layout.

| 1st line | File type, here BOTTLE, followed by a comma and a DATE_TIME stamp YYYYMMDDdivINSwho |
| --- | --- |
|  | YYYY 4 digit year MM 2 digit month DD 2 digit day div division of Institution INS Institution name who initials of responsible person |
| #lines | A file may include 0-N optional lines, typically at the start of a data file, but after the file type line, each beginning with a "#" character and each ending with carriage return or end-of-line. Information relevant to file change/update history of the file itself may be included here, for example. |
| 2nd line | Column headings. |
| 3rd line | Units. |
| data lines | As many data lines may be included in a single file as is convenient for the user, with the proviso that the number and order of parameters, parameter order, headings, units, and commas remain absolutely consistent throughout a single file. |
| END_DATA | The line after the last data line must read END_DATA. |
| other lines | Users may include any information they wish in 0-N optional lines at the end of a data file, after the END_DATA line. |

### Header columns

| Parameter | Format | Description notes |
| --- | --- | --- |
| EXPOCODE | A12 | The expedition code, assigned by the user. |
| SECT_ID | A7 | The SBI station specification. Optional. |
| STNNBR | A6 | The originator's station number. |
| CASTNO | I3 | The originator's cast number. |
| BTLNBR | A7 | The bottle identification number. |
| BTLNBR_FLAG_W | I1 | BTLNBR quality flag. |
| DATE | I8 | Cast date in YYYYMMDD integer format. |
| TIME | I4 | Cast time (UT) as HHMM |
| LATITUDE | F8.4 | Latitude as SDD.dddd where "S" is sign (blank or missing is positive), DD are degrees, and dddd are decimal degrees. Sign is positive in northern hemisphere, negative in southern hemisphere |
| LONGITUDE | F9.4 | Longitude as SDDD.dddd where "S" is sign (blank or missing is positive), DDD are degrees, and dddd are decimal degrees. Sign is positive for "east" longitude, negative for "west" longitude |
| DEPTH | I5 | Reported depth to bottom. Preferred units are "meters" and should be specified in Line 2. In general, corrected depths are preferred to uncorrected depths. Documentation accompanying data includes notes on methodology of correction. Optional. |

### Parameter names, units, and comments:

| Parameter | Units | Format | Comments |
| --- | --- | --- | --- |
| CTDPRS | DB | F9.1 | CTD pressure, decibars |
| CTDPRS_FLAG_W |  | I1 | CTDPRS quality flag |
| SAMPNO |  | A7 | Cast number *100+BTLNBR. Optional |
| CTDTMP | ITS-90 | F9.4 | CTD temperature, degrees C, (ITS-90) |
| CTDTMP_FLAG_W |  | I1 | CTDTMP quality flag |
| CTDCOND | MS/CM | F9.4 | CTD Conductivity, milliSiemens/centimeter |
| CTDCOND_FLAG_W |  | I1 | CTDCOND quality flag |
| CTDSAL |  | F9.4 | CTD salinity |
| CTDSAL_FLAG_W |  | I1 | CTDSAL quality flag |
| SALNTY |  | F9.4 | bottle salinity |
| SALNTY_FLAG_W |  | I1 | SALNTY quality flag |
| SIGMA | THETA | F9.4 | Sigma Theta |
| SIGMA_FLAG_W |  | I1 | Sigma Theta quality flag |
| CTDOXY | UMOL/KG | F9.1 | CTD oxygen, micromoles/kilogram |
| CTDOXY_FLAG_W |  | I1 | CTDOXY quality flag |
| CTDOXY | ML/L | F9.3 | CTD oxygen, milliliters/liter |
| CTDOXY_FLAG_W |  | I1 | CTDOXY quality flag |
| OXYGEN | UMOL/KG | F9.1 | bottle oxygen |
| OXYGEN_FLAG_W |  | I1 | OXYGEN quality flag |
| OXYGEN | ML/L | F9.3 | bottle oxygen, milliliters/liter |
| OXYGEN_FLAG_W |  | I1 | OXYGEN quality flag |
| O2TEMP | DEGC | F6.1 | Temperature of water from spigot during oxygen draw, degrees C |
| O2TEMP_FLAG_W |  | I1 | O2TEMP quality flag |
| SILCAT | UMOL/KG | F9.2 | SILICATE, micromoles/kilogram |
| SILCAT_FLAG_W |  | I1 | SILCAT quality flag |
| SILCAT | UMOL/L | F9.2 | SILCATE, micromoles/liter |
| SILCAT_FLAG_W |  | I1 | SILCAT quality flag |
| NITRAT | UMOL/KG | F9.2 | NITRATE, micromoles/kilogram |
| NITRAT_FLAG_W |  | I1 | NITRAT quality flag |
| NITRAT | UMOL/L | F9.2 | NITRATE, micromoles/liter |
| NITRAT_FLAG_W |  | I1 | NITRAT quality flag |
| NITRIT | UMOL/KG | F9.2 | NITRITE, micromoles/kilogram |
| NITRIT_FLAG_W |  | I1 | NITRIT quality flag |
| NITRIT | UMOL/L | F9.2 | NITRITE, micromoles/liter |
| NITRIT_FLAG_W |  | I1 | NITRIT quality flag |
| PHSPHT | UMOL/KG | F9.2 | PHOSPHATE, micromoles/kilogram |
| PHSPHT_FLAG_W |  | I1 | PHSPHT quality flag |
| PHSPHT | UMOL/L | F9.2 | PHOSPHATE, micromoles/liter |
| PHSPHT_FLAG_W |  | I1 | PHSPHT quality flag |
| NH4 | UMOL/KG | F9.2 | AMMONIUM, micromoles/kilogram |
| NH4_FLAG_W |  | I1 | NH4 quality flag |
| NH4 | UMOL/L | F9.2 | AMMONIUM, micromoles/liter |
| NH4_FLAG_W |  | I1 | NH4 quality flag |
| UREA | UMOL/KG | F9.2 | UREA, micromoles/kilogram |
| UREA_FLAG_W |  | I1 | UREA quality flag |
| UREA | UMOL/L | F9.2 | UREA, micromoles/liter |
| UREA_FLAG_W |  | I1 | UREA quality flag |
| FLUORO | VOLTS | F8.3 | Fluorometer, voltage |
| FLUORO_FLAG_W |  | I1 | Fluorometer quality flag |
| PAR | VOLTS | F8.3 | PAR, voltage |
| PAR_FLAG_W |  | I1 | PAR quality flag |
| SPAR | VOLTS | F8.3 | Surface PAR, voltage |
| SPAR_FLAG_W |  | I1 | Surface PAR quality flag |
| HAARDT | VOLTS | F8.3 | CDOM Fluorometer, voltage |
| HAARDT_FLAG_W |  | I1 | CDOM Fluorometer quality flag |
| N** | UMOL/L | F9.2 | N**, micromoles/liter |
| N**_FLAG_W |  | I1 | N** quality flag |
| CHLORO | UG/L | F8.2 | Chlorophyll, micrograms/liter |
| CHLORO_FLAG_W |  | I1 | Chlorophyll quality flag |
| PHAEO | UG/L | F8.2 | Phaeophytin, micrograms/liter |
| PHAEO_FLAG_W |  | I1 | Phaeophytin quality flag |
| BTL_DEP | METERS | F5.0 | bottle depth, meters |
| BTL_LAT |  | F8.4 | Latitude at time of bottle trip, decimal degrees |
| BTL_LONG |  | F9.4 | Longitude at time of bottle trip, decimal degrees |
| JULIAN |  | F8.4 | Julian day and time as fraction of day of the bottle trip. |

### Quality Flags

CTD data quality flags were assigned to CTDPRS (CTD pressure), CTDTMP (CTD temperature), CTDCOND (CTD Conductivity), and CTDSAL (CTD salinity) as defined in Data Distribution, CTD Data, Quality Flags section of this document. CTDOXY (CTD O2), FLUORO (Fluorometer), PAR (PAR), and SPAR (Surface PAR) parameters are flagged with either a 2, acceptable or 9, not drawn.

Bottle quality flags were assigned to the BTLNBR (bottle number) as defined in the WOCE Operations Manual [Joyce] with the following additional interpretations:

- No problems noted.

- Leaking.  An air leak large enough to produce an observable effect on a sample is identified by a flag of 3 on the bottle and a flag of 4 on the oxygen.  (Small air leaks may have no observable effect, or may only affect gas samples.)

- Did not trip correctly.  Bottles tripped at other than the intended depth were assigned a flag of 4.  There may be no problems with the associated water sample data.

- The samples were not drawn from this bottle.

WHP water sample quality flags were assigned to the water samples using the following criteria:

- The sample for this measurement was drawn from the water bottle, but the results of the analysis were not (yet) received.

- Acceptable measurement.

- Questionable measurement. The data did not fit the station profile or adjacent station comparisons (or possibly CTD data comparisons). No notes from the analyst indicated a problem. The data could be acceptable, but are open to interpretation.

- Bad measurement. The data did not fit the station profile, adjacent stations or CTD data. There were analytical notes indicating a problem, but data values were reported. Sampling and analytical errors were also flagged as 4.

- Not reported. The sample was lost, contaminated or rendered unusable.

- The sample for this measurement was not drawn.

Not all of the quality flags are necessarily used on this data set.

# References

Armstrong, F. A. J., Stearns, C. R., and Strickland, D. H., “The measurement of upwelling and subsequent biological processes by means of the Technicon AutoAnalyzer and associated equipment,” Deep-Sea Research, 14, pp. 381-389, (1967).

Bernhardt, Wilhelms A., “The continuous determination of low level iron, soluble phosphate and total phosphate with the AutoAnalyzer”, Technicon Symposia, I, pp. 385-389 (1967).

Carpenter, J. H., “The Chesapeake Bay Institute technique for the Winkler dissolved oxygen method,” Limnology and Oceanography, 10, pp. 141-143 (1965).

Culberson, C. H., Knapp, G., Stalcup, M., Williams, R.T., and Zemlyak, F., “A comparison of methods for the determination of dissolved oxygen in seawater,” Report WHPO 91-2, WOCE Hydrographic Programme Office (Aug 1991).

Gordon, L.I., Jennings, J.C., Ross, A.A. and J.M. Krest, “A Suggested Protocol for Continuous Flow Automated Analysis of Seawater Nutrients in the WOCE Hydrographic Program and the Joint Global Ocean Fluxes Study”. 1993. WOCE Hydrographic Programs Office, Methods Manual WHPO 91-1.

Gruber, N. and J.L. Sarmiento, “Global patterns of marine nitrogen fixation and denitrification.” Global Biogeochem. Cycles, 11(2), pp. 235-266, (1997).

Intergovernmental Oceanographic Commission, Scientific Committee on Oceanic Research Manual and Guides 29 Protocols for the Joint Global Ocean Flux Study (JGOFS) Core Measurements.  UNESCO, 170pp., (1994).

Joyce, T. ed., and Corry, C. ed., “Requirements for WOCE Hydrographic Programme Data Reporting,” Report WHPO 90-1, WOCE Report No. 67/91 3.1, pp. 52-55, WOCE Hydrographic Programme Office, Woods Hole, MA, USA (May 1994, Rev. 2), UNPUBLISHED MANUSCRIPT

Macdonald, R.W., F. A. McLaughlin and C.S. Wong, “The storage of reactive silicate samples by freezing,” Limnology and Oceanography, 31, pp. 1139-1142 (1986).

Patton, C.J. and Crouch, S.R., “Spectrophotometric and kinetics investigation of the Berthelot reaction for the determination of ammonia,” Analytical Chemistry, 49(3), pp.464-469 (1977).

Rahmatullah, Mohammed, and Boyde, T.R.C, “Improvements in the determination of urea using diacetyl monoxime; methods with and without deproteinisation,” Clinica Chimica Acta, 107, pp.3-9 1980.

Sea-Bird Electronics, Inc, CTD Data Acquisition Software Manual, March 2001.

Sea-Bird Electronics, Inc., CTD Operating and Repair Manual, February 2002.

APPENDIX A: Bottle Quality Comments

Remarks for deleted samples, missing samples, PI data comments, and WOCE codes other than 2 from HLY0403 USCGC Healy. Comments from the Sample Logs and the results of ODF’s investigations are included in this report. Investigation of data may include comparison of bottle salinity and oxygen data with CTD data, review of data plots of the station profile and adjoining stations, and rereading of charts (i.e. nutrients). Units stated in these comments are degrees Celsius for temperature, Practical Salinity Units for salinity, and unless otherwise noted, milliliters per liter for oxygen and micromoles per liter for Silicate, Nitrate, Nitrite, Phosphate and Urea and Ammounium, if appropriate. The first number before the comment is the cast number (CASTNO) times 100 plus the bottle number (BTLNBR).

Station 001.001

101 Sample Log: "Leak from bottom cap after O2 drawn." Oxygen as well as other data

are acceptable.

102 Oxygen and salinity not drawn.

103 Sample Log: "Slight spit from spigot before top vented." Oxygen as well as other

data are acceptable.

104 Sample Log: "Spigot leak when spigot pushed in before venting." Oxygen as well as

other data are acceptable.

105-111 Salinity not drawn.

106 Sample Log: "Slight spigot leak before venting; rotated spigot smaller leak." Oxygen

as well as other data are acceptable.

107 SampleLog: "Oxygen not drawn."

112 Salinity: "Three readings to obtain two good readings." First reading 0.0005 higher

than average, okay as is. PI: "Bottle salinity greater than CTD, would expect the

opposite." Density inversion, 0.01, in the CTD down trace. However the sensor pairs

agree with one another. Leave as is.

Station 002.001

102 SampleLog: "Oxygen and nutrients not drawn."

103 Oxygen: "Overtitrated and backtitrated; no endpoint found."

103-107 Sample Log: "Salinity not drawn."

104 SampleLog: "Spigot leak before venting."

105 SampleLog: "Oxygen redrawn."

106 SampleLog: "Oxygen redrawn."

107 SampleLog: "Spigot leak spigot, before venting."

108 SampleLog: "Salinity, oxygen and nutrients not drawn." Sampled for POM.

109 SampleLog: "Salinity, oxygen and nutrients not drawn." Sampled for C13/N15.

110-111 Sample Log: "Salinity not drawn."

112 PI:"There is a tendency for surface bottle salinity to be a tad higher than CTD salinity."

Comparison of the pair of sensors shows fairly good agreement. Leave as is

Station 003.001

102 Samples taken for POM only, DOM/Lignin only.

104-111 Sample Log: "Salinity not drawn."

108 Samples taken for POM only, DOM/Lignin only.

109 Sample Log: "Bottom cap leak when vented, stopped during sampling." Samples

taken for C13/N15 only.

112 CTD vs. bottle salinity difference large. Primary and secondary sensors also show a

fairly large difference. Ship’s influence apparent in the CTD down trace and even

greater in the up trace. Footnote CTD temperature and salinity questionable.

Station 004.001

101 SampleLog: "Leaking from bottom cap after oxygen drawn." Data are acceptable.

102 Samples taken for POM only.

104 SampleLog: "Spigot leak when vented."

105-107 Salinity not drawn.

107 SampleLog: "Spigot leaking before vented."

108 Samples taken for POM only.

109 Nosamples drawn.

110-111 Salinity not drawn.

Station 005.001

102 Oxygen and salinity not drawn.

103 SampleLog: "Oxygen redraw."

103-105 Salinity not drawn.

104 Oxygen not drawn. Oxygen is acceptable.

106 Sample Log: "Leaking from bottom, bottom cap ajar (o-ring popped out). No samples

taken.

107 SampleLog: "Spigot leak." Data are acceptable.

107-108 Salinity not drawn.

108 See110 CTD comments. Footnote CTD temperature and salinity questionable.

109-110 Oxygen and salinity not drawn.

110 Density inversion at 6 meters. CTD salinity is low compared to bottle. Density inversion

also seen in CTD down trace, probably ship’s effect. Difference, 0.1, between

primary and secondary temperatures. Footnote CTD temperature and salinity questionable.

111 Oxygen not drawn.

Station 006.001

101 SampleLog: "Bottom cap not seated properly-readjusted then okay."

102 No samples drawn per sampling strategy.

104-108 Salinity not drawn.

106 Samples taken for C13/N15 only.

107 No samples drawn per sampling strategy.

109 Sample Log: "Large leak from bottom cap, o-ring popped out." Nutrients are reasonable,

okay as is. Salinity not drawn.

110 Difference, 0.1, between primary and secondary temperatures. Footnote CTD temperature

and salinity questionable. Salinity not drawn.

111 Salinity not drawn.

Station 006.002

201 Sample Log: "Bottom cap leak." Oxygen not drawn per sampling schedule. Nutrients

are acceptable.

202 Salinity, oxygen and nutrients not drawn, sampled for Bacteria, POM and

DOM/Lignin only.

204 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

206 Salinity, oxygen and nutrients not drawn, samples for O2 incubation and taxonomy.

207 Difference between CTD oxygen and bottle could be explained by 1 meters offset

between bottle and probe, strong oxygen gradient.

208 Salinity, oxygen and nutrients not drawn, sampled for POM and DOM/Lignin and

O18.

212 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 2 Salinity not drawn, Productivity cast.

Station 006.003

305 Nutrients drawn per sampling schedule, Radium cast.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 007.001

102 Salinity and oxygen not drawn.

103-108 Salinity not drawn.

104 SampleLog: "Small spigot leak." Oxygen as well as other data are acceptable.

105 CTDLog: "Did not yo-yo, ship rolling some." Data are acceptable.

107 SampleLog: "Small spigot leak." Oxygen as well as other data are acceptable.

109 Salinity and oxygen not drawn.

110 Salinity and oxygen not drawn.

112 SampleLog: "Possible bottom leak." Oxygen as well as other data are acceptable.

Cast 1 Sample Log: "Problem with NaOH dispensor. Lots of jellyfish tentacles." Oxygen

dispensor fixed before cast, data are acceptable.

Station 008.001

102-104 No samples were drawn.

105 Salinity not drawn.

106 Oxygen and salinity were not drawn.

107 Salinity not drawn.

108-109 No samples were drawn.

110 Salinity not drawn.

111 Sample Log: "Air vent found open." Oxygen as well as other data are acceptable.

Salinity not drawn.

112 Large difference between CTD and bottle salinity. Gradient at the surface; could be

the 1 meter difference between bottle and CTD. Primary and secondary sensor agreement,

0.007. Okay as is.

Cast 1 Oxygen temperature probe failure, water in connection and battery low. No oxygen

draw temperature.

Station 008.002

201 Salinity, oxygen and nutrients not drawn, sampled taken for POM and DOM/Lignin.

202 Salinity, oxygen and nutrients not drawn, sampled taken for POM only.

203 Salinity, oxygen and nutrients not drawn, sampled taken for O2 incubation only.

206 Salinity, oxygen and nutrients not drawn, sampled taken for O2 incubation only.

206-212 Biological fouling in primary CTD sensors. CTD oxygen is bad; reporting temperature,

conductivity/salinity from secondary sensors.

208 Salinity, oxygen and nutrients not drawn, sampled taken for POM and DOM/Lignin.

212 Salinity, oxygen and nutrients not drawn, sampled taken for O2 incubation only.

Cast 2 Salinity not drawn, Productivity cast.

Station 009.001

102-106 Salinity not drawn.

104 SampleLog: "Small spigot leak." Data are acceptable.

107 SampleLog: "Bad small spigot leak." Data are acceptable.

Station 010.001

103 SampleLog: "Oxygen redrawn." Oxygen is acceptable.

104 SampleLog: "Small spigot leak."

106-102 Salinity not drawn.

107 SampleLog: "Leaking before venting, sampled from bottle 8."

108 Salinity difference, 0.51, large gradient. Data are acceptable. Primary and secondary

sensor agreement, 0.006. Okay as is.

109-112 Bottles tripped for deck safety, samples not drawn.

Station 011.001

103 SampleLog: "Severe leak coming out of water; no water left for sampling."

104-108 Salinity not drawn.

110-112 Bottle tripped for deck safety, samples not drawn.

Station 012.001

101 SampleLog: "Oxygen flask broken on second shake."

103-109 Salinity not drawn.

106 Sample Log: "Leaking when brought on deck, sampled 7 which was scheduled for

HPLC only."

107 SampleLog: "Small spigot leak." Data are acceptable.

111-112 Duplicate trips, samples not drawn.

Station 013.001

102-106 Salinity not drawn.

107 CTD Log: "Replace bottle prior to cast for small spigot repair." Large difference

between CTD and bottle salinity. Gradient at the surface; could be the 1 meter difference

between bottle and CTD. Primary and secondary sensor agreement, 0.009. Okay

as is.

108-112 Bottles tripped for deck safety, but no samples drawn.

Station 014.001

101-103 Oxygen not drawn per sampling schedule.

103 SampleLog: "Leaking from bottom end cap." Data are acceptable.

104 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation, Bio-optics and

DOM/POM, POM.

105 Sample Log: "Bottom cap leak, after oxygen drawn." Oxygen and other data are

acceptable.

107 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

111 SampleLog: "Oxygen redrawn." Oxygen is acceptable.

112 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 1 Salinity not drawn, Productivity cast.

Station 014.002

201 Nutrients drawn per sampling schedule, POM, and DOM/Lignin.

202 Nutrients drawn per sampling schedule, POM, and DOM/Lignin.

203 Nutrients drawn per sampling schedule, Zooplankton.

205 SampleLog: "Bottom cap leak."

212 Nutrients drawn per sampling schedule, Zooplankton.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 014.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

303 SampleLog: "Bottom cap leak."

305 SampleLog: "Bottom cap leak."

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 015.001

103-111 Salinity not drawn.

104 SampleLog: "Small spigot leak." Oxygen as well as other samples are acceptable.

105 Sample Log: "Bottom cap leak after oxygen drawn, reseated then okay." Oxygen as

well as other samples are acceptable.

111 SampleLog: "Oxygen redrawn." Oxygen rich, data are acceptable.

Station 015.002

201 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 015.003

301 Salinity and oxygen not drawn, sampled for POM.

302 Salinity, oxygen and nutrients not drawn, sampled for DOM/Lignin and O18.

303 Salinity, oxygen and nutrients not drawn, sampled for Bacteria.

304 Salinity and oxygen not drawn, sampled for Bacteria.

305 Salinity and oxygen not drawn, sampled for Bacteria, DOM/Lignin and O18.

307 Salinity, oxygen and nutrients not drawn, sampled for POM, DOM/Lignin and O18.

311 Sample Log: "Bottle cap leak, reseated, then started leaking on venting." Oxygen as

well as other data are acceptable.

Cast 3 Salinity not drawn, Productivity cast.

Station 016.001

103 Salinity and oxygen not drawn per sampling schedule.

104 Salinity not drawn

105 SampleLog: "Bottom cap leak, no oxygen."

105-110 Salinity not drawn.

109 SampleLog: "Small spigot leak." Oxygen and well as other data are acceptable.

112 SampleLog: "Oxygen redrawn. Oxygen and well as other data are acceptable.

Station 017.001

101 CTDLog: "Could not let bottom bottle soak full 2 minutes due to shoaling bottom."

103-111 Salinity not drawn.

104 SampleLog: "Small spigot leak."

105 SampleLog: "Bad bottom leak." Samples not drawn.

110 Density inversion, probably ship effect. Footnote CTD temperature, conductivity/

salinity, and oxygen as questionable.

111 Density inversion, probably ship effect. Footnote CTD temperature, conductivity/

salinity questionable.

Station 018.001

101 CTDLog: "No yo-yo at bottom, could soak only for 1.5 minutes vs. 2 due to shoaling."

Large salinity difference, 0.025, nutrients are acceptable, oxygen is also high

compared with CTD. Leave as is.

103-110 Salinity not drawn.

104 SampleLog: "Small spigot leak." Oxygen as well as other data are acceptable.

Station 019.001

102-111 Salinity not drawn.

Station 020.001

103-111 Salinity not drawn.

Station 021.001

101 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

107 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

Cast 1 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 021.002

201-204 Salinity and oxygen not drawn per sampling schedule.

205 Salinity, oxygen and nutrients not drawn.

206 Oxygen not drawn.

207 Urea lost due to AA malfunction.

Cast 2 Salinity not drawn, Productivity cast.

Station 022.001

Cast 1 Cast aborted, CTD failure.

Station 022.002

201 Sample Log: "Leaked, no samples drawn." Nutrients: "Sample 201 was drawn."

Footnote nutrients bad.

203 SampleLog: "Leaked, no samples drawn."

Station 022.003

Cast 3 Cast aborted, file set up problem.

Station 022.004

504-510 Salinity not drawn.

Station 022.005

504-510 Salinity not drawn.

Station 022.006

601 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

607 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

611 SampleLog: "Leaking from bottom cap, large flow."

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 022.007

701 SampleLog: "Bottom cap leak." Nutrients are acceptable.

701-703 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

705 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

712 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 7 Salinity, oxygen and nutrients not drawn, except as noted, combination Radium and

Bio-Marker cast.

Station 022.008

801-804 Oxygen not drawn per sampling schedule.

803 SampleLog: "Leaking from bottom end cap." Nutrients are acceptable.

804 Nutrients were inadvertently drawn before O2 incubation. Salinity, and oxygen not

drawn, sampled for O2 incubation.

805 Nutrients not drawn from bottle 5, draw was from 4, error in Sample Log.

807 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation.

811 SampleLog: "Leaking from bottom end cap." Bacteria was drawn.

812 Sample Log indicates that sample was taken from bottle, but it was not analyzed.

Footnote nutrients lost.

Cast 8 Salinity not drawn, Productivity cast.

Station 022.009

901 Nutrients drawn per sampling schedule.

9all Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 023.001

104 Sample Log: "Leaking from bottom, no oxygen taken." Bottle is leaking and samples

are bad.

105-111 Salinity not drawn.

Station 023.002

203-212 Salinity not drawn.

Station 023.003

301 SampleLog: "Leaking from bottom cap when came on deck."

303 SampleLog: "Slight leak from bottom cap."

304 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

308 Sample Log: "Leaking from bottom cap." Salinity and oxygen not drawn, nutrients

drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 023.004

409 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 023.005

501-503 Oxygen not drawn and nutrients drawn per sampling schedule.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

505 Analyst entered nutrient sample as 4, but Sample Log says 5. Changed to 5.

508 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 024.001

101 PO4 a little high, corrected peak with bubble, but corrected still seems a tad high.

Footnote PO4 questionable.

102 Sample Log: "Leaking when brought on deck, did not sample for O2." Salinity low

and nutrients bad. Footnote bottle leaking and samples bad.

109-111 Salinity not drawn.

Station 024.002

201-203 Nutrients drawn, samples taken for Bacteria.

204 Salinity, oxygen and nutrients not drawn, sampled for Bacteria.

207 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

212 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 2 Salinity not drawn, Productivity cast.

Station 024.003

303-310 Salinity not drawn.

304 NO3appears low, rechecked data and no problems found.

311 SampleLog: "Leaking from bottom, no O2 drawn." Salinity and nutrients are acceptable.

312 Salinity not drawn.

Station 025.001

101-102 Bottle salinity appears low compared with CTD. Temperature gradient, salinity indicates

that there could have been biological fouling. However, the sensors agree with

one another and similar depth on the next cast. There is no real indication that there is

a problem with the bottle salinities, however bottle 1 was run after a very low salinity.

Footnote bottle salinity questionable.

101-104 Sample Log: "Jellyfish tentacles across bottles."

102 Oxygen not drawn per sampling schedule.

103-110 Salinity not drawn.

104 Sample Log: "Bottom cap leak when spigot open." Oxygen not drawn per sampling

schedule. Nutrients are acceptable.

112 Large CTD-bottle salinity difference, large gradient. Data are acceptable.

Station 025.002

202 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

212 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 025.003

301-303 Nutrients drawn per sampling schedule.

304 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

308 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

312 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 3 Salinity not drawn, Productivity cast.

Station 026.001

101 Sample Log: "Leaking from bottom cap, no oxygen drawn." Salinity not drawn.

Nutrients are acceptable.

102 Oxygen: "Brown precipitate after acidification prevented proper titration-oxygen

lost."

103-104 Salinity not drawn.

104 SampleLog: "Leaking from bottom cap, no oxygen drawn." Nutrients are acceptable.

106-111 Salinity not drawn.

Station 026.002

202 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

211 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 027.001

103-110 Salinity not drawn.

Station 028.001

101 Salinity, oxygen and nutrients not drawn, sampled for POM, DOM/Lignin and O18.

103-104 Nutrients drawn per sampling schedule, Bacteria.

105 Salinity, oxygen and nutrients not drawn, sampled for POM, DOM/Lignin and O18.

109 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 1 Salinity not drawn, Productivity cast.

Station 028.002

201 Nutrients drawn per sampling schedule.

212 Nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 029.001

103 Salinity not drawn.

104 SampleLog: "Bad bottom leak, no samples drawn."

105-110 Salinity not drawn.

Station 029.002

201 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

207 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 030.001

103 SampleLog: "Leaking from bottom cap." No samples drawn.

105 SampleLog: "Leaking from spigot when vented." Data are acceptable.

108 Sample Log: "Leaking from bottom cap." Oxygen and salinity not drawn. Footnote

nutrient data bad.

109 Data are questionable. Not certain what is wrong. Footnote data questionable.

110 Sample Log: "Leaking from bottom cap." Oxygen not drawn. Large CTDS-bottle

difference, two sensors agree with one another. Footnote bottle leaking and samples

bad.

112 SampleLog: "Leaking from bottom cap." No samples drawn, except O18.

Station 030.002

204-211 Salinity not drawn.

205 SampleLog: "Leaking from spigot."

Station 030.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 030.004

402 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 030.005

501-503 Nutrients drawn per sampling schedule, Bacteria.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

508 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 030.006

601 Nutrients drawn per sampling schedule.

612 Nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 030.007

702 Urea appears high, rechecked data and no problem found.

703 Sample Log: "Top o-ring failure." Probably due to lowering bottles after trip due to

winch problem. No samples drawn.

709 Sample Log: "Top cap failure." Probably due to lowering bottles after trip due to

winch problem. No samples drawn.

Cast 7 Salinity and oxygen not drawn, Bio-Marker cast.

Station 030.008

801 Nutrients drawn per sampling schedule, Radium cast.

812 Nutrients drawn per sampling schedule, Radium cast.

Cast 8 Sample Log: "This started out as a test cast to get rid of crossovers on wire spool and

then became a Radium cast." Salinity, oxygen and nutrients not drawn, except as

noted, Radium cast.

Station 031.001

105 SampleLog: "Bottom cap leak, o-ring popped out." No water for sampling.

105-111 Salinity not drawn.

Station 031.002

201 Oxygen appears high. Gradient, oxygen is acceptable.

202-211 Salinity not drawn.

209 SampleLog: "Small spigot leak when vented, top o-ring not sealed." Data are acceptable.

Station 031.003

301-302 Nutrients drawn per sampling schedule, Bacteria.

303 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

306 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

307 No samples taken from this bottle.

309 SampleLog: "Bottom cap leak, could not sample from spigot."

312 Sample Log: "Bottom cap leak." Salinity, oxygen and nutrients not drawn, Bacteria

sampled from this bottle.

Cast 3 Salinity not drawn, Productivity cast.

Station 031.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 031.005

502 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

511 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 032.001

101 SiO3appears high, rechecked data and found no problems.

102 SampleLog: "Large spigot leak, reseated top o-ring then okay."

105-111 Salinity not drawn.

Station 032.002

201-203 Nutrients drawn per sampling schedule, Bacteria.

205 Salinity, oxygen and nutrients not drawn, sampled for C13/N15 only.

211-212 No samples drawn per sampling strategy.

Cast 2 Salinity not drawn, Productivity cast.

Station 032.003

301 Nutrients drawn per sampling schedule.

312 Nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 032.004

403-411 Salinity not drawn.

409 SampleLog: "Air vent open." Data are acceptable.

Station 032.005

503 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

510 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 032.006

609 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

611 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 032.007

701 Nutrients drawn per sampling schedule, Bio-Marker cast.

705 SampleLog: "Top cap did not seal." No samples drawn.

706 Sample Log: "Possible top cap leak. Nutrients drawn per sampling schedule, Bio-

Marker cast. Data are acceptable.

707 Sample Log: "Top cap leak. Water rushed out with spigot open and air vent closed."

No samples drawn.

708-709 Nutrients drawn per sampling schedule, Bio-Marker cast.

711 Nutrients drawn per sampling schedule, Bio-Marker cast.

Cast 7 Salinity, oxygen and nutrients not drawn, except as noted, Bio-Marker cast.

Station 033.001

106-110 Salinity not drawn. Nutrients: "NH4 lost due to ship’s power outage."

111 Sample Log: "Bottom leak." Oxygen not drawn. All samples affected by bottle leak.

Footnote bottle leaking and samples bad.

112 Salinity not drawn.

Cast 1 Oxygen temperature reading approximately 2.5 low, switched to alternate temperature

sensor on bottle 6. Used 0 for 1-5.

Station 033.002

203-210 Salinity not drawn.

206 SampleLog: "Big leak on small spigot." Data are acceptable.

211 SampleLog: "Small spigot leak." Data are acceptable.

Station 033.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 033.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 033.005

501-503 Oxygen not drawn per sampling strategy, Productivity cast.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

506 SampleLog: "Small spigot leak before venting." Data are acceptable.

507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 034.001

102 Sample Log: "Maybe short squirt with NaOH/NaI." Oxygen is acceptable. Salinity

not drawn.

104 SampleLog: "Leaking from bottom when spigot pushed in." Data are acceptable.

104-108 Salinity not drawn.

106 SampleLog: "Small spigot leak." Data are acceptable.

110 Salinity not drawn.

111 SampleLog: "Slight small spigot leak." Data are acceptable.

112 Salinity not drawn.

Cast 1 Sample Log: "Sampling depths modified to allow 3 double bottles on shallow cast for

Bio-Markers."

Station 034.002

201-202 No oxygen drawn.

204 SampleLog: "Small leak from bottom cap." Data are acceptable.

204-209 Salinity not drawn.

206 SampleLog: "Small spigot leak." Data are acceptable.

208 Nooxygen drawn.

211 SampleLog: "Small spigot leak, significant." Data are acceptable.

Station 034.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 034.004

406 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 034.005

501-503 Oxygen not drawn per sampling strategy, Productivity cast.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 034.006

601 Nutrients drawn per sampling schedule.

612 Nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 035.001

101 Oxygen: "Debris in flask, overtitrated, sample lost." Urea seems high, leave as is.

103 Salinity not drawn.

105 Salinity not drawn.

107 SampleLog: "Bottom cap leaks, large flow." Samples not drawn except O18.

108 SampleLog: "Bottom cap leaks, large flow." Samples not drawn except O18.

109-110 Salinity not drawn.

Station 035.002

203-211 Salinity not drawn.

204 Sample Log: "Slow leak on bottom cap, started after O2 draw, reseated then okay."

Oxygen as well as other data are acceptable.

Station 035.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 035.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 035.005

501-502 No samples drawn per sampling strategy.

509-512 No samples drawn per sampling strategy.

Cast 5 Salinity not drawn, Productivity cast.

Station 035.006

Cast 6 Salinity and oxygen not drawn, Bio-Marker cast.

Station 036.001

103-104 CTD vs. bottle salinity high, suspect bottle salinity. Not analytical problems. Agreement

between the two sensors, primary and secondary, is a little off. The primary has

a slight offset at about 1120 meters but then it drifts back. Footnote CTD primary

salinity questionable.

105 Sample Log: "Leaking from bottom cap, slow leak." Comparison between O2 probe

and bottle oxygen looks okay.

105-110 Salinity not drawn.

109 Sample Log: "Leaking from bottom cap, reseated then slight leak, reventing started

the leak again." Samples not drawn except O18.

Station 036.002

202-211 Salinity not drawn.

Station 036.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 036.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 036.005

501-503 Oxygen not drawn per sampling strategy, Productivity cast.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 036.006

601-602 Nutrients drawn per sampling schedule.

611-612 Nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 037.001

105-110 Salinity not drawn.

106 SampleLog: "Minor spigot leak." Data are acceptable.

111 SampleLog: "Top cap canted significant small spigot leak." Data are acceptable.

Station 037.002

202-205 Salinity not drawn.

207-210 Salinity not drawn.

212 Salinity not drawn.

Station 037.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 037.004

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 037.005

501-502 Oxygen not drawn per sampling strategy, Bacteria.

503 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

506-507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

508 Oxygen appears low compared with CTD. No analytical problem noted. Oxygen is

acceptable.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 037.006

605 SampleLog: "Bad bottom cap leak." No samples drawn.

607 SampleLog: "Bad bottom cap leak." No samples drawn.

608 SampleLog: "Good size small spigot leak." Nutrients are acceptable.

Cast 6 Salinity and oxygen not drawn, Bio-Marker cast.

Station 038.001

102 Oxygen appears low, strong gradient. Data are acceptable.

103-104 Salinity, oxygen and nutrients not drawn, sampled for C13/N15 only.

109-112 No samples drawn per sampling strategy.

Cast 1 Salinity not drawn, Productivity cast.

Station 038.002

201 All values suspect, poor flushing or leaking bottle. Footnote bottle leaking, data questionable.

203 SampleLog: "Bottom cap leak, small." Data are acceptable.

203-211 Salinity not drawn.

205 Duplicate trip, no samples drawn.

208 Duplicate trip, no samples drawn.

210 Duplicate trip, no samples drawn.

Station 038.003

301-303 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

303 Rechecked Nitrate could not find any analytical problem. Footnote NO3 questionable.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule. Rechecked

Nitrate could not find any analytical problem. Footnote NO3 questionable.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 039.001

102 Duplicate trip samples not drawn.

104 Duplicate trip samples not drawn.

106 Duplicate trip samples not drawn.

108 Duplicate trip samples not drawn.

110 Duplicate trip samples not drawn.

111 SampleLog: "Spigot leak, sampling moved to 12." Duplicate trip samples not drawn.

111-102 Salinity not drawn.

Station 040.001

102 Oxygen not drawn.

102-111 Salinity not drawn.

104 Oxygen not drawn.

106 Oxygen not drawn.

108 Oxygen not drawn.

110 Oxygen not drawn.

111 SampleLog: "Small spigot leak." Data are acceptable.

112 Oxygen not drawn.

Station 041.001

101 SampleLog: "Leaked from bottom cap." Duplicate level, samples not drawn.

103-111 Salinity not drawn.

104 Salinity, oxygen and nutrients not drawn, sampled for thorium.

106 Salinity, oxygen and nutrients not drawn, sampled for thorium.

107 SampleLog: "Bottom cap leak with vent open-check o-ring." Data are acceptable.

108 Salinity, oxygen and nutrients not drawn, sampled for thorium.

110 Oxygen not drawn.

111 SampleLog: "Small spigot leak." Data are acceptable.

112 Oxygen not drawn.

Station 042.001

103-111 Salinity not drawn.

107 Salinity, oxygen and nutrients not drawn, duplicate bottle trip, sampled for C13/N15

only.

Station 042.002

201 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

212 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 042.003

301-302 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

303 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

305 Oxygen not drawn.

306 Oxygen low, okay strong gradient, interleaving.

307 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

308 Oxygen high, okay strong gradient, interleaving.

312 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 3 Salinity not drawn, Productivity cast.

Station 042.004

401 Nutrients drawn per sampling schedule.

412 Nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 043.001

102 Oxygen not drawn.

103-105 Salinity not drawn.

104 Oxygen not drawn.

106 Oxygen not drawn.

108 Oxygen not drawn.

108-111 Salinity not drawn.

110 Oxygen not drawn.

111 SampleLog: "Small spigot leak." Data are acceptable.

112 Oxygen and salinity not drawn.

Station 044.001

101 SampleLog: "Bottom cap leak." Salinity, oxygen and nutrients not drawn.

102 Salinity not drawn.

104 Salinity not drawn.

105 SampleLog: "Top cap leak." Salinity, oxygen and nutrients not drawn.

106 Salinity not drawn.

109-111 Salinity not drawn.

111 Oxygenl ow, okay large gradient.

Station 044.002

201-202 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

207-208 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

212 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Radium cast.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 044.003

301-303 Nutrients drawn per sampling schedule.

304 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

307 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

312 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 3 Salinity not drawn, Productivity cast.

Station 045.001

102 Oxygen not drawn.

102-111 Salinity not drawn.

104 Oxygen not drawn.

106 Oxygen not drawn.

107 SampleLog: "Small spigot leak, minor." Data are acceptable.

108 Oxygen not drawn.

110 Oxygen not drawn.

111 SampleLog: "Small spigot leak." Data are acceptable.

112 Oxygen not drawn.

Station 046.001

101 SampleLog: "Bottom cap leak." No samples drawn.

103-111 Salinity not drawn.

104 Oxygen not drawn.

106 Oxygen not drawn.

108 Oxygen not drawn.

110 Oxygen not drawn.

112 Oxygen not drawn.

Station 047.001

102 Oxygen not drawn.

102-105 Salinity not drawn.

108-111 Salinity not drawn.

110 SampleLog: "Small spigot leak, miniscule." Data are acceptable.

111 Sample Log: "Small spigot leak." Rechecked NO3 and found no problems.

Rechecked CTDO, appears to be reasonable, gradient area. Data are acceptable.

Station 047.002

202-204 Salinity not drawn.

206-211 Salinity not drawn.

210 SampleLog: "Small spigot leak." Data are acceptable.

211 SampleLog: "Small spigot leak." Data are acceptable.

Station 047.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

303 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

305 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

307 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

309 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

311 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Bio-Marker cast.

Station 047.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 047.005

501 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 047.006

601-603 Oxygen not drawn per sampling strategy.

604 No samples drawn per sampling strategy.

605 SampleLog: "Small spigot leak." Data are acceptable.

607 No samples drawn per sampling strategy.

609 Sample Log: "Small spigot leak when vented, top cap not seated." Data are acceptable.

612 No samples drawn per sampling strategy.

Cast 6 Salinity not drawn, Productivity cast.

Station 048.001

106-111 Salinity not drawn.

108 SampleLog: "Spigot leak." Data are acceptable.

109 SampleLog: "Top cap did not close properly-no samples drawn."

110 SampleLog: "Top cap did not close properly-no samples drawn."

111 SampleLog: "Spigot leak, questionable sample." Data are acceptable.

Station 048.002

203 SampleLog: "Slight bottom cap leak." Data are acceptable.

204-208 Salinity not drawn.

208 SampleLog: "Slight spigot leak." Data are acceptable.

210 SampleLog: "Slight spigot leak." Data are acceptable.

211 SampleLog: "Spigot leak." Data are acceptable.

212 Salinity not drawn.

Station 048.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

311 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 048.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 048.005

501-503 Oxygen not drawn per sampling strategy.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

510-512 CTD salinity at duplicate trip +/-0.004 different from each other. Salinity at 10

meters could be a little low. Suspect temperature effect. Within accuracy leave as is.

Ship’s effect could be the cause, but usually ship’s effect gives higher temperature.

511 SampleLog: "Small spigot leak." Data are acceptable.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 048.006

601 Nutrients drawn per sampling schedule.

612 Nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 049.001

Cast 1 Cast aborted, bottle 11 bottom cap closed when rosette entered the water.

Station 049.002

206 SampleLog: "Very small spigot leak." Data are acceptable.

206-211 Salinity not drawn.

208 SampleLog: "Spigot leak, bad." Data are acceptable.

210 SampleLog: "Spigot leak." Data are acceptable.

211 SampleLog: "Small spigot leak." Data are acceptable.

Station 049.003

301-302 Salinity not drawn.

304-306 Salinity not drawn.

308 SampleLog: "Bad spigot leak." Data are acceptable.

310 Sample Log: "Spigot leak." Salinity: "Replaced due to repeated failure to hold thimble."

Salinity as well as other data are acceptable.

312 Salinity not drawn.

Station 049.004

401 SampleLog: "Bottom cap leak." Data are acceptable.

401-403 Oxygen not drawn per sampling strategy.

405 Oxygen not drawn per sampling strategy.

406 SampleLog: "Small and big spigot leak, top o-ring popped." Data are acceptable.

408-409 CTDO questionable, slight offset when bottle tripped.

410 Sample Log: "Bottom cap leak." No samples drawn per sampling strategy, duplicate

trip.

410-412 There is a large difference in the temperature for the 3 bottles that were tripped at the

surface. Leave as is.

412 No samples drawn per sampling strategy, duplicate trip.

Cast 4 Salinity not drawn, Productivity cast.

Station 049.005

501 Salinity and oxygen not drawn, nutrients drawn per sampling schedule. Rechecked

urea peaks and data are acceptable.

512 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 049.006

601 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 049.007

701 Sample Log: "Bottom cap leak." Salinity, oxygen and nutrients not drawn. Coded

bottle leaking.

702-707 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

705 Sample Log: "Top cap leak." Nutrient duplicate with bottle 4 agree very well. Data

are acceptable.

708 Sample Log: "Bottom cap leak." Salinity, oxygen and nutrients not drawn. Coded

bottle leaking.

709-712 Nutrients drawn per sampling schedule.

710 SampleLog: "Spigot leak." Data are acceptable.

Cast 7 Salinity, oxygen and nutrients not drawn, except as noted, Bio-Marker cast.

Station 050.001

104-105 Salinity not drawn.

107 Salinity not drawn.

108 SampleLog: "Spigot leak." Data are acceptable.

109 SampleLog: "Spigot leak, slight." Salinity not drawn. Data are acceptable.

110 SampleLog: "Spigot leak." Data are acceptable.

111 Salinity not drawn.

112 Salinity: "Replaced bottle due to repeated failure to hold thimble." Salinity appears

high compared with CTD. Footnote salinity bad.

Station 050.002

206 SampleLog: "Spigot leak from top cap, skewed." Data are acceptable.

210 Salinity not drawn.

212 Salinity: "Three readings to obtain two good readings." Salinity is acceptable. Bottle

oxygen appears a little high, gradient. Data rechecked and no problem found with the

analyses.

Station 050.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 050.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 050.005

501-503 Oxygen not drawn per sampling schedule.

504 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

507 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

509 Sample Log: "Bottom cap leak, slight, okay when reseated, noticed after oxygen

draw." Data are acceptable.

512 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 5 Salinity not drawn, Productivity cast.

Station 051.001

108 SampleLog: "Significant small spigot leak." Data are acceptable.

110 SampleLog: "Significant small spigot leak." Data are acceptable.

Cast 1 Salinity and oxygen not drawn, nutrients drawn per sampling schedule, Bio-Marker

cast.

Station 051.002

202-205 Salinity not drawn.

206 Sample Log: "Minor spigot leak." Data are acceptable. CTD salinity if a little low.

Primary and secondary comparison indicate that primary is approximately 0.007 low.

Within accuracy, leave as is.

206-212 CTD oxygen bad.

208 SampleLog: "Significant spigot leak." Data are acceptable.

210 SampleLog: "Just a spigot leak." Data are acceptable.

211-212 Salinity not drawn.

Station 051.003

302-311 Salinity not drawn.

Station 051.004

401-402 Oxygen not drawn per sampling strategy.

403 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

406 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

409 SampleLog: "Small spigot leak." Data are acceptable.

411 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

412 No samples drawn per sampling strategy.

Cast 4 Salinity not drawn, Productivity cast.

Station 051.005

501 Nutrients drawn per sampling schedule.

512 Nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 051.006

601 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

603 SampleLog: "Bottom cap leak."

612 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 051.007

702 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

711 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 7 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 052.001

102-104 Salinity not drawn.

106 Salinity not drawn.

108 Salinity not drawn.

110-111 Salinity not drawn.

Station 052.002

203 Sample Log: "Leaking from bottom cap, reseated then okay." No samples drawn.

Code bottle leaking.

204 Sample Log: "Leaking from bottom cap, reseated then okay." No oxygen drawn.

Data are not acceptable. Salinities show bad agreement, other data are all wrong.

Footnote bottle leaking and samples bad.

206 SampleLog: "Small spigot leak." Data are acceptable.

208 SampleLog: "Small spigot leak." Data are acceptable.

208-211 Salinity not drawn.

Station 052.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 052.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 053.001

103-105 Salinity not drawn.

105 Sample Log: "Bottom cap leak, small leak, reseated then okay." Data are acceptable.

106 SampleLog: "Small spigot leak." Data are acceptable.

107-109 Salinity not drawn.

112 Salinity not drawn.

Station 053.002

204 Salinity not drawn.

206 SampleLog: "Slow small spigot leak." Data are acceptable.

206-211 Salinity not drawn.

208 SampleLog: "Slow small spigot leak." Data are acceptable.

Station 054.001

103-111 Salinity not drawn.

106 SampleLog: "Slow small spigot leak."

Station 054.002

201-203 Nutrients drawn per sampling schedule, Bacteria. Oxygen not drawn per sampling

strategy.

204 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

206 SampleLog: "Spigot leaks, both." Data are acceptable.

207 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

212 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

Cast 2 Salinity not drawn, Productivity cast.

Station 054.003

303-305 Salinity not drawn.

307 Salinity not drawn.

308 Sample Log: "Spigot leak." Salinity: "Salinity bottle filled too full, above neck."

Data are acceptable.

309 Salinity not drawn.

310 Sample Log: "Spigot leak." Salinity: "Salinity bottle filled too full, above neck."

Data are acceptable.

312 Salinity not drawn.

Station 054.004

406 SampleLog: "Bottom cap leak, when vent open." Data are acceptable.

408 SampleLog: "Minor spigot leak." Data are acceptable.

410 SampleLog: "Bottom cap leak." No samples drawn.

Cast 4 Salinity and oxygen not drawn, Bio-Marker cast.

Station 054.005

501 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

512 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 054.006

601 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

612 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 6 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 055.001

102-106 Salinity not drawn.

103 SampleLog: "Bad bottom cap leak." No samples drawn.

108 SampleLog: "Spigot leak." Data are acceptable.

109 Salinity not drawn.

110 SampleLog: "Spigot leak."

111 SampleLog: "Possible top cap leak." Data are acceptable.

112 Salinity not drawn.

Station 055.002

201-202 Salinity not drawn.

203 SampleLog: "Small bottom cap leak." Data are acceptable

204-206 Salinity not drawn.

208 SampleLog: "Spigot leak." Data are acceptable

209 Salinity not drawn.

210 Sample Log: "Minor spigot leak." Salinity: "Won’t hold thimble, replaced." Data are

acceptable

211-212 Salinity not drawn.

Station 055.003

308 SampleLog: "Spigot leak." Data are acceptable.

Cast 3 Salinity and oxygen not drawn, Bio-Marker cast.

Station 055.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 055.005

507 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 055.006

601 Sample Log: "Leaking from bottom cap." Salinity, oxygen and nutrients not drawn.

Bacteria drew from bottom cap.

602-603 Salinity and oxygen not drawn per sampling strategy.

606 Salinity, oxygen and nutrients not drawn, sampled for O2 incubation only.

611-612 No samples drawn per sampling strategy.

Cast 6 Salinity not drawn, Productivity cast.

Station 055.007

701 Nutrients drawn per sampling schedule.

712 Nutrients drawn per sampling schedule.

Cast 7 Salinity, oxygen and nutrients not drawn, except as noted, Zooplankton cast.

Station 056.001

101-105 Salinity not drawn.

104 SampleLog: "Top cap leak." No samples drawn.

106 SampleLog: "Minor bottom cap leak." Data are acceptable.

108 SampleLog: "Spigot leak." Data are acceptable.

110 SampleLog: "Minor spigot leak." Data are acceptable.

111-112 Salinity not drawn.

Station 056.002

201 Sample Log: "Bottom cap leak." Nutrients appear to have been drawn from bottle 2.

Footnote nutrients questionable.

202-204 Salinity not drawn.

205-206 Sample Log: "Slight bottom cap leak." Data are acceptable.

207 Salinity not drawn.

208 SampleLog: "Spigot leak." Data are acceptable.

209 Salinity not drawn.

210 SampleLog: "Spigot leak." Data are acceptable.

211-212 Salinity not drawn.

Station 056.003

303 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

305 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

306 Sample Log: "Bottom cap leak after venting." Bio-Markers switched all sampling to

bottle 5. Nutrients drawn. Data are acceptable. Salinity and oxygen not drawn, nutrients

drawn per sampling schedule.

309 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Bio-Marker cast.

Station 056.004

401 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

412 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 4 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 056.005

501 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 5 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 057.001

101 SampleLog: "Leaking from bottom cap, reseated then okay." Data are acceptable.

101-103 Salinity and oxygen not drawn per sampling strategy.

105 Salinity and oxygen not drawn per sampling strategy, duplicate trip, nutrients only.

111-112 No samples drawn per sampling strategy.

Cast 1 Salinity not drawn, Productivity cast.

Station 058.001

103 Rechecked urea peak and there was no analytical problem.

103-105 Salinity not drawn.

106 SampleLog: "Slight bottom cap leak after spigot was opened." Data are acceptable.

107 Salinity not drawn.

109 Salinity not drawn.

110 SampleLog: "Slight spigot leak." Data are acceptable.

111 Salinity not drawn.

Station 058.002

201 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

212 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 2 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 059.001

101 SampleLog: "Bottom cap leak on venting, reseated then okay." Data are acceptable.

103-111 Salinity not drawn.

106 Sample Log: "Bottom cap leak after oxygen draw, reseated then okay." Data are

acceptable.

109 SampleLog: "Slight spigot leak." Data are acceptable.

Station 059.002

201 Salinity and oxygen not drawn per sampling strategy.

202 Nutrients sampled before oxygen. Data are acceptable.

203 Salinity and oxygen not drawn per sampling strategy.

204 No samples drawn per sampling strategy, duplicate trip.

208 No samples drawn per sampling strategy, duplicate trip.

210 No samples drawn per sampling strategy, duplicate trip.

212 No samples drawn per sampling strategy, duplicate trip.

Cast 2 Salinity not drawn, Productivity cast.

Station 059.003

301 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

312 Salinity and oxygen not drawn, nutrients drawn per sampling schedule.

Cast 3 Salinity, oxygen and nutrients not drawn, except as noted, Radium cast.

Station 060.001

101 SampleLog: "Bottom cap leak." Data are acceptable.

102-107 Salinity not drawn.

106 SampleLog: "Slight bottom cap leak." Data are acceptable.

109 SampleLog: "Slight bottom cap leak." Data are acceptable. Salinity not drawn.

111 Salinity not drawn.

Station 060.002

Cast 2 Salinity, oxygen and nutrients not drawn, Radium cast. Bottles were sampled and

drained before nutrients were drawn.