IMOS underway CO2 dataset report

Dataset:  AA_2014_029-059_V4 fCO2 IMOS.txt

Voyage Information:

MV Aurora Australis voyage AA1314V4

Departed: 	Hobart, Tasmania, January 30, 2014

Arrived:  	Hobart, Tasmania, March 01, 2014

Principal Investigator:

Bronte Tilbrook, CSIRO Oceans and Atmosphere

PO Box 1538, Hobart, Tasmania 7001 Australia

e: bronte.tilbrook@csiro.au

Track:

Ship Details:

Name: Aurora Australis

Call Sign: VNAA

Home Port: Hobart, Australia

Ship owner: P&O Polar

Citation

Publications that use these data should reference the data source as:

Tilbrook, B., C. Neill and J. Akl (2014) Integrated Marine Observing System underway CO2 data for Aurora Australis voyage AA1314V4, Integrated Marine Observing System

CO2 System Overview:

The fugacity of carbon dioxide (fCO2) in surface seawater was measured using a General Oceanics Inc. automated system (Model 8050; Pierrot et al 2009). Seawater is sprayed into an equilibration chamber and CO2 in the headspace gas equilibrates with the seawater. The headspace gas is pumped through a thermoelectric condenser followed by a nafion drying tube, before flowing through a Licor 7000 non-dispersive infrared gas analyser used to measure the CO2 mole fraction (XCO2) of the dried air. The gas flow is stopped temporarily for the CO2 measurements, which are made at atmospheric pressure. A set of four CO2 standards (Table 1) that cover the range of CO2 values expected in the ocean are analysed about every three hours to calibrate the gas analyser. Atmospheric XCO2 (dry) is measured after the standards by pumping clean outside air from an intake on the forward mast of the ship.

Table 1. CO2-in-air standard values measured on the WMO-X2007 mole fraction scale

| Cylinder no. | Cylinder number | CO2 (ppm) |
| --- | --- | --- |
| 1 | CA06898 | 0.0 |
| 2 | CA01610 | 299.41 |
| 3 | CA01669 | 353.01 |
| 4 | CA01673 | 402.15 |

Seawater intake and ancillary data

The seawater intake is at about 6 m depth. A remote temperature sensor (Seabird Electronics SBE38) located at the intake is used to measure sea surface temperature (SST). Sea surface salinity is measured using a thermosalinograph (Seabird Electronics SBE21) mounted in the oceanographic lab next to the pCO2 system. The travel time between the intake and CO2 system is typically about 70 seconds with warming usually less than 0.7ºC in cold waters near freezing. The thermosalinograph water is from the same intake and supply line. Meteorological data, salinity, SST, and ships position and time are taken from the ship’s logging system.

System Description:

Equilibrator:  			Showerhead, water volume of ~0.5 L, headspace ~ 0.8 L.

Water flow rate:  		2 to 3.0 l/min

Headspace gas flow rate:  	60-120 ml/min

Measurement method:  		Infrared absorption of dried gas.

CO2 Sensor: 			Licor 7000, Serial # IRG4-0910

Resolution/Uncertainty: 		±0.3 atm for equilibrator and atmospheric fCO2

Equilibrator temperature: 	Hart 1521 thermometer (S/N A8B266), 5610-9 probe (S/N B0727114), accuracy ± 0.01°C.

Equilibrator pressure: 		Setra Model 239 (S/N 2223344), accuracy ± 0.15 hPa.

Atmospheric pressure:	Druck RPT350 (S/N 2729757), accuracy ± 0.15 hPa. calibrated annually against Druck DPI141 barometric sensor and Bureau of Meteorology barometric pressure standard.

Sea Surface Temperature:  	Seabird SBE38, accuracy ± 0.01°C.

Salinity: 	Seabird SBE 21 thermosalinograph mounted in the   oceanographic lab next to the pCO2 system, accuracy ±0.01.

Data Fields and Units:

| Field | Name | Unit | Description |
| --- | --- | --- | --- |
| 1. | Group/ship | - | CSIRO/Astrolabe |
| 2. | CruiseID | AAYyYyV# | Cruise designation |
| 3. | JD_GMT | ddd.hhhh | Decimal day and time of year, GMT time |
| 4. | Date | yyyymmdd | 20111231 |
| 5. | Time | hh:mm:ss | UTC time |
| 6. | Lat | degrees | Latitude, decimal degrees |
| 7. | Long | degrees | Longitude, decimal degrees |
| 8. | xCO2EQ_PPM | ppm | Mole fraction of CO2 in the equilibrator head space (dry). |
| 9. | xCO2ATM_PPM | ppm | Mole fraction of CO2 in the atmosphere (dry) measured every 4 hours after standard runs |
| 10. | xCO2ATM_PPM _INTERPOLATED | ppm | Mole fraction of CO2 in the atmosphere (dry) measured every 4 hours after standard runs and values linearly interpolated to the times shown |
| 11. | Press_Equil | hPa | Equilibrator head space pressure |
| 12. | Press_ATM | hPa | Barometric pressure |
| 13. | TEQ | ºC | Equilibrator water temperature |
| 14. | SST | ºC | Sea surface temperature |
| 15. | SAL | psu | Sea surface salinity |
| 16. | fCO2SW_UATM | atm | Fugacity of carbon dioxide at surface water salinity and temperature |
| 17. | fCO2ATM_UATM _INTERPOLATED | atm | fugacity of CO2 in the atmosphere |
| 18. | DfCO2 | atm | fCO2SW - fCO2ATM |
| 19. | LICORflow | ml/min | Gas flow through infrared gas analyser |
| 20. | H2Oflow | lpm | Water flow to equilibrator |
| 21. | WindSpd_True | m/s | Wind speed. |
| 22. | WindDirn_True | degrees | Wind direction, 0 is North and 90 is East. |
| 23. | Type | - | Measurement type (equilibrator, standard or atmosphere) |
| 24. | WOCE_QC_FLAG | - | 2 = Good 3 = Questionable 4 = Bad (data identified as bad are not reported). |
| 25. | SUBFLAG | - | Secondary flags, only for questionable measurements, WOCE flag 3 (Pierrot et al 2009): - Outside of standard range - Questionable/interpolated SST - Questionable EQU temperature - Anomalous (EQU T-SST)(+0.6ºC) - Questionable sea-surface salinity - Questionable pressure - Low EQU gas flow - Questionable air value - Other, water flow |

Quality control and data reduction:

Parameters logged by the fCO2 system and ship sensors are quality controlled after each voyage.

- Data with missing parameters or obvious outliers for the ship or fCO2 system parameters are marked as missing and removed from the calculations. Parameter values are flagged as good (flag=2), questionable (flag=3), or bad (flag=4), depending on the range of values expected. Many of the ship and CO2 system parameters are not reported in the final dataset, but are used to establish that the system is functioning correctly. For example, water flow rates to the equilibrator below 2 LPM are flagged as questionable and the cause investigated with the flag value changed to 4 if the flow has been interrupted or is insufficient. Similar checks are made to ensure the gas flow through the infrared gas analyser is in a suitable range (50 to 120ml/min). The list of parameters checked are:

CO2 system data quality controlled:

GPS date and Time

Latitude and Longitude

Water flow rate

Gas flow rates through licor analyser

Atmospheric pressure

Equilibrator pressure

Equilibrator water temperature

Mole fraction CO2

Water vapour in gas stream

Licor NDIR temperature

Ship’s data quality controlled:

GPS date and time

Latitude and Longitude

Sea surface temperature

Sea surface salinity

Relative wind speed and direction

True wind speed and direction

- The data sets are next evaluated for excessive warming of the seawater flowing to the equilibrator, and for contamination of the atmospheric measurements by ship stack gas.

The fCO2 value in the water is sensitive to warming between the ship intake and equilibrator. The travel time between the ship intake and equilibrator is first checked by comparing the timing of rapid changes in surface water temperature for the intake (SST) and the equilibrator temperatures. The travel time or lag time is normally about 70 seconds. The warming in the system used on MV Aurora Australis is typically about 0.4 °C, increasing to about 0.6°C in cooler regions.

Atmospheric CO2 values can be influenced by contamination from ship stack gas. The atmospheric air intake is located on the Port side above the wheelhouse to collect air in the front part of the ship within about 20m of the ship stacks. The relative wind speed and direction recorded by the ship meteorological sensors are used to if anomalous atmospheric measurements could be due to stack gas contamination. Data where wind speeds are above 3ms-1 and with a direction of ±60° of the bow are typically good values. Data with likely stack gas contamination are flagged as bad (flag = 4) and not included in the calculations outlined below.

- After completion of the quality control checks, the measured mole fractions are corrected to final values using measurements of the four CO2-in-air standards (Table 1). The standards are run about every four hours to bracket the air and equilibrator measurements. The offsets between the measured and certified values of each standard are linearly interpolated to the times of measurement of the air and equilibrator samples. At each measurement time, a linear regression of offset values versus certified standard values is used to calculate the offset to apply to the measured air and equilibrator values. The corrections are typically small (less than 1 ppm) and account for drift of the gas analyser response over time. The corrected mole fractions (dry) for the equilibrator and air samples flagged as good are then used to calculate the fugacity of CO2. Only data flagged as good or suspect are reported in the final data set.

fCO2SW and fCO2ATM:

The fugacity of carbon dioxide in seawater is determined using the following equation (Weiss, 1974; Dickson et al, 2007):

fCO2eq =  XCO2(Peq - pH2O)exp(Peq(B + 2δ)/(R•Teq))

where XCO2 is the mole fraction (dry) in the equilibrator headspace, P is the pressure in the equilibrator; pH2O is the water vapour pressure (Weiss and Price, 1980) at the temperature of water in the equilibrator (Teq) and its salinity:

pH2O(atm) = exp(24.4543 - 67.4509(100/Teq) - 4.8489ln(Teq/100) - 0.000544S)

R the ideal gas constant (82.0578 cm3·atm/K·mol), B the virial coefficient of pure CO2, and the cross virial coefficient of a CO2-air mixture (Weiss, 1974).

B(cm3/mol) = -1636.75 + 12.0408Teq – 0.032795Teq2 + 0.0000316528Teq3

δ(cm3/mol) =  57.7 – 0.118Teq

An empirical correction (Copin-Montegut, 1988) is applied to account for warming of water between the sea surface and equilibrator. The same equations are applied to the measurements of the mole fraction of CO2 in atmospheric gas, using the sea surface temperature and atmospheric pressure.

The air-sea gradient in fCO2 is calculated as: DfCO2 = fCO2SW - fCO2ATM

where fCO2SW is the bulk surface seawater value and fCO2ATM the atmospheric value.

Processing Comments:

The ship’s underway thermosalinograph (salinity), sea surface temperature and meteorological data were collected and calibrated by the Australian Antarctic Division Data Centre.

While the ship was in heavy ice between 17/02/14 14:47:48 and 21/02/14 10:39:48 the salinity records were affected by melting ice inside the de-bubbler. These values were flagged as ‘bad’ and excluded for calculations.

The time lag between the intake temperature and the equilibrator temperature on this cruise is of less than 70 seconds. This time lag was not applied to the SST data.

Acknowledgements:

SOOP-CO2 data was sourced as part of the Integrated Marine Observing System (IMOS) – supported by the Australian Government through the National Collaborative Research Infrastructure Strategy and the Super Science Initiative. Marie Bigot helped ran the pCO2 system during this cruise.

References

Copin-Montegut, C. (1988) A new formula for the effect of temperature on the partial pressure of CO2 in sea water, Marine Chemistry, 25, p29-37 (incl. Corrigedum, Marine Chemistry (1989) 27, pp143-144).

Dickson, A.G., C. Sabine and J. R. Christian (2007) Guide to best practices for Ocean CO2 measurements. PICES Special Publ. 3, 191 pp.

Pierrot, D., C. Neill, K. Sullivan, R. Castle, R. Wanninkhof, H. Lüger, T. Johannessen, A. Olsen, R. A. Feely, C. E. Cosca (2009) Recommendations for Autonomous Underway pCO2 Measuring Systems and Data Reduction Routines, Deep-Sea Research II, doi:10.1016/j.dsr2.2008.12.005

Weiss, R. F (1974) Carbon Dioxide in water and sea water: the solubility of a non-ideal gas, Marine Chemistry, 2, pp.203-215

Weiss, R.F. and B. A. Price (1980) Nitrous oxide solubility in water and seawater. Marine Chemistry 8, 347–359.