## metadata
Instructions: Please try not to change the order of Rows No. 1 through No. 211, as the information will be read by a computer computer program later on. Starting from No. 212, please first append the additional variable sections, then the non-measured variable sections, then the additional principal investigator sections (if there are more than three PIs), and then the platform sections (if there are more than 3 platforms). Please do not use special characters.
No | Metadata element name | Your input | Help reference no.
1 | Submission Date | 2017-03-20 | 1
2 | Accession no. of related data sets |  | 2
3 | Investigator-1 name | Yuichiro Takeshita | 3.1
4 | Investigator-1 institution | Monterey Bay Aquarium Research Institute | 3.2
5 | Investigator-1 address | 7700 Sandholdt Road, Moss Landing, CA 95039 | 3.3
6 | Investigator-1 phone | 831-775-1713 | 3.4
7 | Investigator-1 email | yui@mbari.org | 3.5
8 | Investigator-1 researcher ID | 0000-0003-1824-8517 | 3.6
9 | Investigator-1 ID type (ORCID, Researcher ID, etc.) | ORCID | 3.7
10 | Investigator-2 name | Todd R. Martz | 3.1
11 | Investigator-2 institution | Scripps Institution of Oceanography | 3.2
12 | Investigator-2 address | 9500 Gilman Drive, La Jolla, CA 95093 | 3.3
13 | Investigator-2 phone | 858-534-7466 | 3.4
14 | Investigator-2 email | trmartz@ucsd.edu | 3.5
15 | Investigator-2 researcher ID | 0000-0002-1996-7145 | 3.6
16 | Investigator-2 ID type (ORCID, Researcher ID, etc.) | ORCID | 3.7
17 | Data submitter name | Yuichiro Takeshita | 4.1
18 | Data submitter institution | Monterey Bay Aquarium Research Institute | 4.2
19 | Data submitter address | 7700 Sandholdt Road, Moss Landing, CA 95039 | 4.3
20 | Data submitter phone | 831-775-1713 | 4.4
21 | Data submitter email | yui@mbari.org | 4.5
22 | Data submitter researcher ID | 0000-0003-1824-8517 | 4.6
23 | Data submitter ID type (ORCID, Researcher ID, etc.) | ORCID | 4.7
24 | Title | Measurements of pH (on total scale) and other variables collected from surface underway observations using Durafet pH electrode and Chloride Ion Selective Electrode and other instruments from NOAA Ship R/V Ronald H. Brown in the North Atlantic Ocean and South Atlantic Ocean during the CLIVAR/GO-SHIP Repeat Section A13.5_2010 (EXPOCODE 33RO20100308) from 2010-03-08 to 2010-04-17 (NCEI Accession 0162231) | 5
25 | Abstract | An automated underway pH system was operated in the hydro lab of the R/V Ronald H. Brown during the CLIVAR/GO-SHIP Repeat Section A13.5 cruise in 2010. pH was measured every 30 seconds using a Durafet III pH combination electrode and an Orion Chloride Ion Selective Electrode (Cl-ISE). This combination allows for the calculation of pH using two reference electrodes: the internal reference electrode of the Durafet (pHint) and the Cl-ISE directly submersed in seawater (pHext). pH is reported on 30 second intervals to provide the maximum flexibility for assessing precision and capturing small scale features in the ocean. Seawater continually flowed through a flowcell that housed the Durafet and Cl-ISE, then a SBE45 Micro Thermosalinograph. The temperature inside the flowcell was assumed to be the same as the SBE45. The salinity measurements from the SBE45 were not reliable on this cruise (based on comparisons with discrete bottle samples), therefore salinity from the SBE45 located next to the underway fCO2 system [Wanninkhof, 2012] was interpolated onto the pH time stamp. Latitude, longitude, and sea surface temperature (SST) were also interpolated from the underway fCO2 data set. All pH is reported on the total scale, and at SST. All data reported here are considered “good”. We recommend the use of pHext, as we believe this is more accurate than pHint. | 6
26 | Purpose | Mapping surface pH | 7
27 | Start date | 2010-03-08 | 8.1
28 | End date | 2010-04-17 | 8.2
29 | Westbd longitude | -3.7 | 9.1
30 | Eastbd longitude | 9.7 | 9.2
31 | Northbd latitude | 4.9 | 9.3
32 | Southbd latitude | -54.4 | 9.4
33 | Geographic names | Eastern part of the South Atlantic | 11
34 | Funding agency name |  | 13.1
35 | Funding project title |  | 13.2
36 | Funding project ID (Grant no.) |  | 13.3
37 | Research projects | (GO-SHIP) CLIVAR-CO2 | 14
38 | Platform-1 name | Ronald H. Brown | 15.1
39 | Platform-1 ID | 33RO | 15.2
40 | Platform-1 type | Research Vessel | 15.3
41 | Platform-1 owner | NOAA | 15.4
42 | Platform-1 country | USA | 15.5
43 | Platform-2 name |  | 15.1
44 | Platform-2 ID |  | 15.2
45 | Platform-2 type |  | 15.3
46 | Platform-2 owner |  | 15.4
47 | Platform-2 country |  | 15.5
48 | Platform-3 name |  | 15.1
49 | Platform-3 ID |  | 15.2
50 | Platform-3 type |  | 15.3
51 | Platform-3 owner |  | 15.4
52 | Platform-3 country |  | 15.5
53 | EXPOCODE | 33RO20100308 | 16
54 | Cruise ID | RB0907-Leg2 | 17
55 | Section | A13.5 | 18
56 | Author list for citation | Takeshita, Yuichiro; Martz, Todd, R; | 19
57 | References | Takeshita, Y., Johnson, K. S., Martz, T. R., and Sarmiento, J. (2017) Assessment of autonomous pH measurements for determining surface seawater partial pressure of CO2, Geophysical Research Letters, in review. | 20
58 | Supplemental information | See attached word document for further details on calibration of the pH data | 21
59 | DIC: Variable abbreviation in data files |  | 22.1
60 | DIC: Variable unit |  | 22.5
61 | DIC: Observation type |  | 22.2
62 | DIC: Measured or calculated |  | 22.6
63 | DIC: Calculation method and parameters |  | 22.7
64 | DIC: Sampling instrument |  | 22.8
65 | DIC: Analyzing instrument |  | 22.9
66 | DIC: Detailed sampling and analyzing information |  | 22.10
67 | DIC: Field replicate information |  | 22.11
68 | DIC: Standardization technique description |  | 22.12.1
69 | DIC: Frequency of standardization |  | 22.12.2
70 | DIC: CRM manufacturer |  | 22.12.3.1
71 | DIC: Batch number |  | 22.12.3.2
72 | DIC: Poison used to kill the sample |  | 22.13.1
73 | DIC: Poison volume |  | 22.13.2
74 | DIC Poisoning correction description |  | 22.13.3
75 | DIC: Uncertainty |  | 22.14
76 | DIC: Data quality flag description |  | 22.15
77 | DIC: Method reference (citation) |  | 22.16
78 | DIC: Researcher Name |  | 22.17.1
79 | DIC: Researcher Institution |  | 22.17.2
80 | TA: Variable abbreviation in data files |  | 23.1
81 | TA: Variable unit |  | 23.5
82 | TA: Observation type |  | 23.2
83 | TA: Measured or calculated |  | 23.6
84 | TA: Calculation method and parameters |  | 23.7
85 | TA: Sampling instrument |  | 23.8
86 | TA: Analyzing instrument |  | 23.9
87 | TA: Type of titration |  | 23.1
88 | TA: Cell type (open or closed) |  | 23.11
89 | TA: Curve fitting method |  | 23.12
90 | TA: Detailed sampling and analyzing information |  | 23.13
91 | TA: Field replicate information |  | 23.14
92 | TA: Standardization technique description |  | 23.15.1
93 | TA: Frequency of standardization |  | 23.15.2
94 | TA: CRM manufacturer |  | 23.15.3.1
95 | TA: Batch Number |  | 23.15.3.2
96 | TA: Poison used to kill the sample |  | 23.16.1
97 | TA: Poison volume |  | 23.16.2
98 | TA: Poisoning correction description |  | 23.16.3
99 | TA: Magnitude of blank correction |  | 23.17
100 | TA: Uncertainty |  | 23.18
101 | TA: Data quality flag description |  | 23.19
102 | TA: Method reference (citation) |  | 23.2
103 | TA: Researcher Name |  | 23.21.1
104 | TA: Researcher Institution |  | 23.21.2
105 | pHA: Variable abbreviation in data files | pH_EXT_TOT_AT_SST
106 | pHA: pH scale | Total
107 | pHA: Observation type | Surface Underway
108 | pHA: Measured or calculated | Measured
109 | pHA: Calculation method and parameters
110 | pHA: Sampling instrument | Ship Underway system
111 | pHA: Analyzing instrument | Honeywell Durafet III pH electrode and Orion Chloride Ion Selective Electrode
112 | pH: Temperature of measurement | in-situ | 24.1
112 | pHA: Detailed sampling and analyzing information | Please refer to the separately uploaded methods description file. Details on the sensor theory and calculations can be found in [Martz et al., 2010; Bresnahan et al., 2014], but will be briefly described here. Computation of pHext requires knowledge of two sensor calibration coefficients, the sensor reference potential at 0 °C 〖(E〗_(ext,0°C)^*) and the temperature coefficient ((dE_ext^*)/dT): 〖pH〗_ext= ((E_ext-E_ext^* )+S×log⁡(γ_H γ_Cl m_Cl))/S (1) E_ext^*=E_(ext,0°C)^*+(dE_ext^*)/dT T (2) where Eext is the electromotive force between the ISFET and the Cl-ISE, E_ext^* is the reference potential at in situ temperature, T is temperature in Celsius, γ is the activity coefficient of the respective ions, mCl is the molality of chloride, and S = R×TK×ln(10)/F (R is the universal gas constant 8.3145 J K-1 mol-1, TK is temperature in Kelvin, and F is the Faraday constant 96485 C mol-1). S, γ_H γ_Cl, and m_Cl are calculated from measurements of temperature and salinity [Khoo et al., 1977; Dickson et al., 2007; Martz et al., 2010]. A linear regression between E_ext^* and T gives E_(ext,0°C)^* (intercept) and (dE_ext^*)/dT (slope). (dE_ext^*)/dT was determined to be -957 ± 8 (95% CI) µV °C-1. Computation of pHint requires knowledge of two sensor calibration coefficients, the sensor reference potential at 0 °C 〖(E〗_(int,0°C)^*) and the temperature coefficient ((dE_int^*)/dT): 〖pH〗_int= ((E_int-E_int^* ))/S (3) E_int^*=E_(int,0°C)^*+(dE_int^*)/dT T (4) where Eint is the electromotive force between the ISFET and the internal reference electrode, E_int^* is the reference potential at in situ temperature, A linear regression between E_int^* and T gives E_(int,0°C)^* (intercept) and (dE_int^*)/dT (slope). (dE_int^*)/dT was determined to be -1.127 ± 0.008 (95% CI) mV °C-1. Both E_int^* and E_ext^* can be calculated by measuring Eint or Eext in a solution of known pH. Here, we used the surface discrete sample for fCO2 and dissolved inorganic carbo (DIC) from the rosette at each station to calculate the calibration coefficients (n = 129). E_(ext,0°C)^* was determined to be -1.37370 ± 0.0003 (1σ) V. A non-linear drift in E_(int,0°C)^* was observed, with a range of approximately 2 mV (Figure 1). In order to account for this drift, a piece-wise linear regression was performed over three sections: 1) before March 21st, 2) between March 21 and 28, and 3) after march 28. pHint was calculated using this interpolated E_(int,0°C)^*. The cause of this drift is not known, but could be largely corrected for due to the large number of calibration samples. Figure 1: Time series of E_(int,0°C)^* (open circles). Piece wise linear regression is shown in the solid black line. All carbonate calculations were done using CO2SYS [van Heuven et al., 2011], using carbonate equilibrium constants from [Mehrbach et al., 1973] refit by [Dickson and Millero, 1987], salinity to boron ratios from [Uppstrom, 1974], and bisulfate equilibrium constants from [Dickson, 1990]. We believe the underway pHext is accurate to 0.007, based on the consistency of calibrating the pH sensors using multiple combinations of discrete samples (i.e. fCO2, total alkalinity, DIC, and spectrophotometric pH). However, the short term precision (hours) and long term precision (weeks) of the underway pHext data is 0.0005 and 0.004, respectively. Details of the uncertainties associated with pHext can be found in [Takeshita et al., 2017, in review]. Variable Descriptions DATE_UTC_ddmmyyyy: Date in UTC, formatted as ddmmyyyy TIME_UTC_hhmmss: Date in UTC, formatted as hh:mm:ss LAT_dec_degree: Latitude in decimal degree, interpolated from the underway fCO2 system LONG_dec_degree: Longitude in decimal degree, interpolated from the underway fCO2 system Eint_volt: Voltage between the ISFET and internal reference electrode Eext_volt: Voltage between the ISFET and the external reference electrode (Cl-ISE) TEMP_FLOWCELL_C: Temperature of the flow cell in Celsius SALINITY_PSS_78: Sea surface salinity based on the PSS-78 practical salinity scale,, interpolated from the underway fCO2 system SEA_SURFACE_TEMPERATURE_C: Sea surface temperature in Celsius, interpolated from the underway fCO2 system pH_INT_TOT_AT_SST: pH on the total scale at sea surface temperature calculated using the internal reference electrode on the total scale, at SST pH_EXT_TOT_AT_SST: pH on the total scale at sea surface temperature calculated using the external reference electrode on the total scale, at SST References Bresnahan, P. J., T. R. Martz, Y. Takeshita, K. S. Johnson, and M. LaShomb (2014), Best practices for autonomous measurement of seawater pH with the Honeywell Durafet, Methods Oceanogr., 9, 44–60. Dickson, A. G. (1990), Standard potential of the reaction : AgCl ( s ) + 1/2H2( g ) = Ag ( s ) + HCl ( aq ), and the standard acidity constant of the ion HSO4- in synthetic sea water from 273 . 15 to 318 . 15 K, J. Chem. Thermodyn., 22(3), 113–127. Dickson, A. G., and F. J. Millero (1987), A Comparison of the Equilibrium Constants for the Dissociation of Carbonic Acid in Seawater Media, Deep Sea Res., 34(10), 1733–1743. Dickson, A. G., C. L. Sabine, and J. R. Christian (Eds.) (2007), Guide to best practices for ocean CO2 measurements, PICES Special Publication 3. van Heuven, S., D. Pierrot, J. W. B. Rae, E. Lewis, and D. W. . Wallace (2011), MATLAB Program Developed for CO2 System Calculations. ORNL/CDIAC-105b., , doi:10.3334/CDIAC/otg.CO2SYS_MATLAB_v1.1. Khoo, K. H., R. W. Ramette, C. H. Culberson, and R. G. Bates (1977), Determination of hydrogen ion concentrations in seawater from 5 to 40.degree.C: standard potentials at salinities from 20 to 45%, Anal. Chem., 49(1), 29–34, doi:10.1021/ac50009a016. Martz, T. R., J. G. Connery, and K. S. Johnson (2010), Testing the Honeywell Durafet for seawater pH applications, Limnol. Oceanogr. Methods, 8, 172–184, doi:10.4319/lom.2010.8.172. Mehrbach, C., C. H. Culberson, J. E. Hawley, and R. M. Pytkowicz (1973), Measurement of the apparent dissociation constants of carbonic acid in seawater at atmospheric pressure, Limnol. Oceanogr., 18(6), 897–907, doi:10.4319/lo.1973.18.6.0897. Uppstrom, L. R. (1974), The boron/chlorinity ratio of deep-sea water from the Pacific Ocean, Deep. Res. Oceanogr. Abstr., 21(2), 161–162, doi:10.1016/0011-7471(74)90074-6. Wanninkhof, R.H. (2012). Underway measurements of surface partial pressure of CO2 during the R/V Ronald H. Brown Cruise in the Atlantic Ocean on CLIVAR Repeat Hydrography Section A13.5_2010 (March 8 - April 17, 2010). http://cdiac.ornl.gov/ftp/oceans/CLIVAR/A13.5_2010.data/Underway_data/. Carbon Dioxide Information Analysis Center, Oak Ridge National Laboratory, US Department of Energy, Oak Ridge, Tennessee. doi: 10.3334/CDIAC/OTG.CLIVAR_A13.5_2010_UNDERWAY
113 | pHA: How was temperature inside the flowcell measured | SBE45 directly in line with flow cell
114 | pHA: How was salinity measured | SBE45 from the underway pCO2 was used
115 | pHA: ISFET Driving Electornics Version | Honeywell Cap Adapter
116 | pHA: Durafet Model Number and Serial Number | N/A
117 | pHA: Standardization technique description | pH data were calibrated using surface bottle measurements at each station (DIC, TA, spectrophotometric pH, and fCO2). An average calibration coefficient (k0) is used for the entire dataset (n = 129 for calibraiton samples)
118 | pHA: Frequency of standardization | A single calibration coefficient was used for the cruise. Calibration samples were taken approximately every 6 hours
119 | pHA: E*ext(0C) | (-)1.3737 V
120 | pHA: dE*ext/dT | (-)0.957 mV C-1
121 | pHA: E*int(0C) | Drift corrected; see attached document for details
122 | pHA: dE*int/dT | (-)1.127 mV C-1
123 | pHA: Temperature correction method | CO2SYS, using pH and TA (estimated from sea surface temperautre and salinity) as inputs, and equilibrium constants from [Mehrbach et al. 1973] refit by [Dickson and Millero, 1987],
124 | pHA: at what temperature was pH reported | SST
125 | pHA: Uncertainty | Accuracy: 0.007; Short term precision (hours): 0.0005; Long term precision (weeks): 0.004
126 | pHA: Data quality flag description | All data considered good
127 | pHA: Method reference (citation) | Martz, T. R., J. G. Connery, and K. S. Johnson (2010), Testing the Honeywell Durafet for seawater pH applications, Limnol. Oceanogr. Methods, 8, 172–184, doi:10.4319/lom.2010.8.172.
128 | pHA: Researcher Name | Yuichiro Takeshita
129 | pHA: Researcher Institution | Monterey Bay Aquarium Research Institute
130 | pCO2A: Variable abbreviation in data files |  | 25.1
131 | pCO2A: Variable unit |  | 25.5
132 | pCO2A: Observation type |  | 25.2
133 | pCO2A: Measured or calculated |  | 25.6
134 | pCO2A: Calculation method and parameters |  | 25.7
135 | pCO2A: Sampling instrument |  | 25.8
136 | pCO2A: Location of seawater intake |  | 25.9
137 | pCO2A: Depth of seawater intake |  | 25.1
138 | pCO2A: Analyzing instrument |  | 25.11
139 | pCO2A: Detailed sampling and analyzing information |  | 25.12
140 | pCO2A: Equilbrator type |  | 25.13.1
141 | pCO2A: Equilibrator volume (L) |  | 25.13.2
142 | pCO2A: Vented or not |  | 25.13.3
143 | pCO2A: Water flow rate (L/min) |  | 25.13.4
144 | pCO2A: Headspace gas flow rate (L/min) |  | 25.13.5
145 | pCO2A: How was temperature inside the equilibrator measured . |  | 25.13.6
146 | pCO2A: How was pressure inside the equilibrator measured. |  | 25.13.7
147 | pCO2A: Drying method for CO2 gas |  | 25.14
148 | pCO2A: Manufacturer of the gas detector |  | 25.15.1
149 | pCO2A: Model of the gas detector |  | 25.15.2
150 | pCO2A: Resolution of the gas detector |  | 25.15.3
151 | pCO2A: Uncertainty of the gas detector |  | 25.15.4
152 | pCO2A: Standardization technique description |  | 25.16.1
153 | pCO2A: Frequency of standardization |  | 25.16.2
154 | pCO2A: Manufacturer of standard gas |  | 25.16.3.1
155 | pCO2A: Concentrations of standard gas |  | 25.16.3.2
156 | pCO2A: Uncertainties of standard gas |  | 25.16.3.3
157 | pCO2A: Water vapor correction method |  | 25.17
158 | pCO2A: Temperature correction method |  | 25.18
159 | pCO2A: at what temperature was pCO2 reported |  | 25.19
160 | pCO2A: Uncertainty |  | 25.2
161 | pCO2A: Data quality flag description |  | 25.21
162 | pCO2A: Method reference (citation) |  | 25.22
163 | pCO2A: Researcher Name |  | 25.23.1
164 | pCO2A: Researcher Institution |  | 25.23.2
165 | pCO2D: Variable abbreviation in data files |  | 26.1
166 | pCO2D: Variable unit |  | 26.5
167 | pCO2D: Observation type |  | 26.2
168 | pCO2D: Measured or calculated |  | 26.6
169 | pCO2D: Calculation method and parameters |  | 26.7
170 | pCO2D: Sampling instrument |  | 26.8
171 | pCO2D: Analyzing instrument |  | 26.9
172 | pCO2D: Storage method |  | 26.1
173 | pCO2D: Seawater volume (mL) |  | 26.11
174 | pCO2D: Headspace volume (mL) |  | 26.12
175 | pCO2D: Temperature of measurement |  | 26.13
176 | pCO2D: Detailed sampling and analyzing information |  | 26.14
177 | pCO2D: Field replicate information |  | 26.15
178 | pCO2D: Manufacturer of the gas detector |  | 26.16.1
179 | pCO2D: Model of the gas detector |  | 26.16.2
180 | pCO2D: Resolution of the gas detector |  | 26.16.3
181 | pCO2D: Uncertainty of the gas detector |  | 26.16.4
182 | pCO2D: Standardization technique description |  | 26.17.1
183 | pCO2D: Frequency of standardization |  | 26.17.2
184 | pCO2D: Temperature of standardization |  | 26.17.3
185 | pCO2D: Manufacturer of standard gas |  | 26.17.4.1
186 | pCO2D: Concentrations of standard gas |  | 26.17.4.2
187 | pCO2D: Uncertainties of standard gas |  | 26.17.4.3
188 | pCO2D: Water vapor correction method |  | 26.18
189 | pCO2D: Temperature correction method |  | 26.19
190 | pCO2D: at what temperature was pCO2 reported |  | 26.2
191 | pCO2D: Uncertainty |  | 26.21
192 | pCO2D: Data quality flag description |  | 26.22
193 | pCO2D: Method reference (citation) |  | 26.23
194 | pCO2D: Researcher Name |  | 26.24.1
195 | pCO2D: Researcher Institution |  | 26.24.2
196 | Var1: Variable abbreviation in data files | SEA_SURFACE_TEMPERATURE_C | 27.1
198 | Var1: Full variable name | Sea Surface Temperature | 27.2
197 | Var1: Variable unit | ITS-90 | 27.7
199 | Var1: Observation type |  | 27.4
200 | Var1: Sampling instrument | Ship's Underway system | 27.1
201 | Var1: Analyzing instrument | SBE58 | 27.11
202 | Var1: Detailed sampling and analyzing information | The SST temperature measurements are from the Seabird 21 TSG, which is immersed in the instrument chest of the ship in the bow thruster space. SST data were interpolated from the underway fCO2 dataset to the timestamp of the pH sensor | 27.13
203 | Var1: Field replicate information | N/A | 27.14
204 | Var1: Uncertainty | Factory calibraiton better than 0.01 C; not verified on cruise | 27.15
205 | Var1: Data quality flag description |  | 27.16
206 | Var1: Method reference (citation) |  | 27.17
207 | Var1: Researcher Name | Rik Wanninkhof | 27.21.1
208 | Var1: Researcher Institution | NOAA/AOML | 27.21.2
209 | Var2: Variable abbreviation in data files | SALINITY_PSS_78 | 27.1
211 | Var2: Full variable name | Practical Salinity | 27.2
210 | Var2: Variable unit | PSS 78 | 27.7
212 | Var2: Observation type |  | 27.4
213 | Var2: Sampling instrument | Ship's Underway system | 27.1
214 | Var2: Analyzing instrument | SBE45 | 27.11
215 | Var2: Detailed sampling and analyzing information | The salinity measurements are from the Seabird 45 Micro TSG in the underway fCO2 system. Salinity data were interpolated from the underway fCO2 dataset to the timestamp of the pH sensor | 27.13
216 | Var2: Field replicate information | N/A | 27.14
217 | Var2: Uncertainty | Better than 0.1, based on comparisons to surface bottle salinity measurements (n = 129) | 27.15
218 | Var2: Data quality flag description | All data considered good | 27.16
219 | Var2: Method reference (citation) |  | 27.17
220 | Var2: Researcher Name | Rik Wanninkhof | 27.21.1
221 | Var2: Researcher Institution | NOAA/AOML | 27.21.2