Palmer Leg 97-03; April-May 1997; Hydrographic data methodology

US JGOFS Antarctic Environments Southern Ocean Process Study (AESOPS)

Palmer Leg 97-03; April-May 1997

Documentation for: THE HYDROGRAPHIC BOTTLE DATA

L.A. Codispoti (lou@ccpo.odu.edu)

Old Dominion University, May 1998

General Comments:

This "readme" file pertains to the salinity, dissolved oxygen, and

nutrient data taken from sampling bottles with the hydrographic

rosette that was equipped with 24 ~10-liter "Niskin-like" Bullister

Bottles made mostly of PVC and equipped with orange silicone o-rings

during Palmer leg 97-03 (April-May 1997). Dr. Hugh Ducklow of the

Virginia Institute of Marine Sciences (duck@vims.edu) was the chief

scientist during this leg. This cruise was the third process study leg

of the U.S. JGOFS program in the Southern Ocean (AESOPS). Several

casts with a Trace Metal clean rosette equipped with 8, 30-l Go-Flo

bottles were also taken during this leg, but they are not reported

here because this system was not designed to produce hydrographic data

of "WOCE quality".

Some questionable data are not included in this report. These data

together with data from the trace metal casts are retained in files at

Old Dominion University and are available upon request.

No units are given for salinity in this report because the most recent

definitions of salinity define it as a dimensionless number.  To

accommodate every preference, Winkler oxygen values are reported in

ml/l, micromolar and micromoles per kg.  The latter values can only be

calculated with a knowledge of the oxygen sample temperatures when the

samples were drawn. These "draw temperatures" are not reported here,

but can be obtained by contacting lou@ccpo.odu.edu.  Nutrient values

are reported in micromolar. They can be converted to micromoles per

kg, by combining laboratory temperature on the Palmer (approx. 21

deg C during this leg) and the salinity of the sample to compute

density and then dividing the value in micromolar by this number.

Methods:

In general, the methods employed for the bottle salinity, Winkler

dissolved oxygen, and nutrient analyses did not differ significantly

from those described in the JGOFS protocols that were distributed in

1994 (UNESCO, IOC Manual and Guide #29). Minor differences included

the following: 1) Sea Bird CTD systems and bottle carousels were

employed (SBE-9+ underwater units, SBE-11 deck units, SBE-32

carousels).  These units represent a newer generation of equipment

than the units described in the JGOFS protocols.  2) The weights of

the salts used for primary standards for dissolved oxygen were not

adjusted to an "in vacuo" basis as suggested in the protocols. It is

unlikely that this departure from procedure would cause significant

errors. Our calculations suggest that the maximum differences arising

from our decision to not correct to an "in vacuo" basis would be 0.02%

3) The protocols give one a choice of adjusting nutrient methods so

that calibration curves are strictly linear, or opting for more

response and taking into account non-linearities.  We choose the

former method.  4) No corrections were made for "carryover" between

nutrient samples run on the Technicon Autoanalyzer. Carryover effects

in our nutrient analyses are generally less than ~2% of the

concentration difference between adjacent samples, and were minimized

by arranging samples in depth order and by running duplicate samples

in some cases.  5) Calibration and re-calibration of volumetric ware

were not exactly as described in the JGOFS protocols.  6)Duplicate

oxygen samples were not routinely drawn. 7) The JGOFS protocols do not

describe an automated technique for the analysis of ammonium

concentrations.  We employed the Berthelot reaction using a

modification of the method described by Whitledge et al. (1981,

Whitledge, T.E., Malloy, S.C., Patton, C.J. and Wirick, C.D. Automated

Nutrient Analyses in Seawater. Brookhaven National Laboratory Rept.

BNL 51398, 216pp.).

Temperature:

The temperature data associated with each bottle depth were taken by

the CTD system during the bottle tripping process.  Consult the

companion CTD data report for this cruise to learn more about the CTD

system.

Sampling:

The samples in this report were taken from ~10 liter Bullister

bottles. NOTE THAT THE BULLISTER BOTTLE IN THE NUMBER 5 POSITION

(NB10B025) ON THE HYDROGRAPHIC ROSETTE WAS FOUND TO HAVE BLACK SPIGOT

O-RINGS UP TO AND INCLUDING STATION 16 CAST 6.  THIS SHOULD HAVE NO

EFFECT ON THE HYDROGRAPHIC DATA REPORTED HERE, BUT ANY BIOLOGICIAL

INCUBATION SAMPLES TAKEN FROM THIS BOTTLE COULD BE SUSPECT.

Because there is little or no lag time between triggering a bottle and

bottle closure with the new SeaBird rosette systems, our sampling

protocols request that bottles be held at the sampling depth for at

least 30 seconds before tripping.

NOTE THAT THE MID-POINT OF THE SAMPLING BOTTLES WAS 0.8 METER ABOVE

THE CTD SENSORS.  THE DATA HAVE NOT BEEN CORRECTED FOR THIS OFFSET.

MORE IMPORTANTLY, THE CONDITIONS ON THIS CRUISE OFTEN PREVENTED

COMPLETE FLUSHING OF THE BULLISTER BOTTLES DESPITE THE REQUESTED 30

SECOND SOAK TIME. EXAMINATION OF BOTTLE VS CTD SALINITIES IN GRADIENTS

SUGGEST THAT THE BOTTLES OFTEN RETAINED SOME DEEPER WATER WHEN THEY

WERE TRIPPED AND HAD APPARENT DEPTHS A FEW METERS DEEPER THAN THE CTD

PROBE.  IN GENERAL, THE SALINITY DIFFERENCES COULD BE EXPLAINED BY

DEPTH OFFSETS OF 5 METERS OR LESS FROM THE CTD PROBE ASSUMING THAT ALL

OF THE WATER IN THE BOTTLE CAME FROM THE DEEPER DEPTH. IN FACT, THE

BOTTLE WOULD HAVE CONTAINED A BLEND OF WATER FROM DEPTHS ABOVE AND

BELOW THE CTD SENSOR BUT WITH AN APPARENT BIAS TOWARDS THE DEEPER

DEPTHS.  THERE WERE A FEW APPARENT OFFSETS OF ABOUT 10 METERS BUT THE

DIFFERENCES IN APPARENT DEPTH WERE SELDOM GREATER THAN 5 METERS AND IN

MOST CASES SIGNIFICANTLY LESS THAN 5 METERS. BY COMPARING BOTTLE AND

CTD SALINITY DATA IN HIGH GRADIENT REGIONS, THE USER CAN ASSESS THE

IMPORTANCE OF THIS EFFECT FOR A PARTICULAR CAST.

Salinity:

Salinities were determined with Guildline Autosal salinometers. New

vials of standard sea-water were used to standardize before and at the

end of every run. Agreement between bottle salinities and the recently

calibrated sensors on the Sea Bird CTD systems was usually better than

0.02 (except in regions of strong gradients) before post-cruise data

processing which employs the bottle salinities to correct the CTD

salinities. More information on the quality of the salinity data are

given in the companion CTD report, but we note that the initial

salinity data from the primary CTD sensors drifted by about 0.01

during this cruise. Both the corrected CTD salinity data collected at

the time of bottle tripping and the salinities run on the Niskin bottle

samples with an Autosal salinometer are reported here.

Dissolved oxygen:

The Winkler dissolved oxygen set-up was built and supplied by the

SIO/ODF group.  This system is computer controlled and detects the

end-point photometrically.  Temperatures of the thiosulfate and

standard solutions are automatically monitored by this system.

Nutrients:

Note that the terminology used to describe nutrients has become

somewhat loose over the years and that silicate=silicic acid,

dissolved silicon or reactive silicate, and phosphate=reactive

phosphorus.

Nutrient analyses were performed on a 5-channel Technicon II AA system

that was modified and provided by Dr. Lou Gordon of Oregon State

University.

During PALMER leg 96-04a, the nutrient standards provided by

Dr. Gordon's group were compared with standards from the Ocean Data

Facility Group at Scripps and with standards purhased from Ocean

Scientific International (OSI). Interested parties may contact

Dr. Louis I. Gordon (lgordon@oce.orst.edu) if they are interested in

the details of these intercomparisons.  The only notable differences

were a tendency for the OSI silicate signals to be ~1% high relative

to the OSU standards,and 4% low relative to OSU nitrite standards.  We

believe that the OSI nitrite standards were in error. NOTE THAT DURING

THIS LEG (PALMER 97-03), MAJOR PROBLEMS WERE ENCOUNTERED WITH NITRITE

STDS. WE HAVE CORRECTED THE NITRITE DATA TO AN ACCURACY OF ABOUT 10%

DURING POST-CRUISE DATA PROCESSING. SINCE WATER COLUMN NITRITE VALUES

WERE GENERALLY LESS THAN ~0.25 MICROMOLAR MAXIMUM REMAINING ERRORS IN

THE WATER COLUMN NITRITE DATA ARISING FROM THIS PROBLEM SHOULD BE ON

THE ORDER OF 0.02 TO 0.03 MICROMOLAR. Nitrate concentrations have been

adjusted for the changes in nitrite values since nitrate is obtained

by subtracting nitrite from the results of the nitrate+nitrite

analysis. Because of the concentration of effort over the Ross Sea

shelf and the season, significant amounts of nitrite and ammonium could

occur throughout the water column, and the usual method of looking at

deep-water values to correct nitrite and ammonium baselines could not

be employed.  Thus, there may be some baseline uncertainties in nitrite

and ammonium concentrations that would induce errors of less than ~0.1

micromolar for ammonium and less than ~0.02 micromolar for nitrite.

Tests of the salt effect in the ammonium analysis during this cruise

suggested that it was negligible for the water column values reported

here, but this may not be the case for specialized incubation and ice

samples. Most Cd column efficiencies appeared to be >95%, but

correction factors were applied to the data from stations 06 casts 03,

04, and 06 (Cd column efficiency = 0.7), and from station 19 cast 01

(Cd column efficiency = 0.86) because of low Cd column efficiencies and

the use of mixed working standards. These standards contained between

30 to 31 micromolar nitrate and between 1 to 2 micromolar nitrite

depending on the state of degradation of thenitrite standards which

were being oxidized to nitrate during this cruise (the nitrite problem

alluded to above).  Water column nitrite values were so low (generally

less than 0.2 micromolar) that it was not found necessary to correct

for Cd column efficiency when subtracting nitrite values from the

nitrate plus nitrite analyses to obtain nitrate concentratrions.

Queries:

Questions about these data may be addressed to:

Dr. L.A. Codispoti

CCPO

Old Dominion University

Norfolk, VA 23529

lou@ccpo.odu.edu