Walker O. Smith

University of Tennessee

POC-PON Method:

Sampling

Samples are collected in 1 or 2 liter polypropylene bottles.  Water is

filtered onto precombusted (450 degrees C, 3 hours) 25mm Whatman GF/F

filters (nominal pour size 0.7 micron).  Just prior to end of

filtration, filter is rinsed with a weak HCl (0.01 N in filtered sea

water) to remove inorganic carbonates.  Filter is removed, placed in a

precombusted (525 degrees C, 2 hours) shell vial, and vial is capped

with precombusted (525 degrees C, 2 hours) aluminum foil.  Samples are

stored in a 60 degree C oven until processed

Procedure

Sample Analysis

1.      Prior to analysis, the filters are kept dry in a 60 degree

oven, in precombusted (525 degree C, 2 hours) vials.

2.      The samples are placed in tin capsules and analyzed on a Carla

Erba EA1108-Elemental Analyzer.  One conditioning sample, two tin blank

samples, three standard samples, and three standard samples run as

unknowns are run to test the machines calibration each day. An average

of twenty to twenty-five POC/PON samples can be analyzed each day.  A

standard is run after approximately five POC/PON samples for

calibration.

Standardization and blank determination

Acetanilide standards and tin blanks (empty tin sleeves) are measured

prior to each batch of samples. The standard is weighed in tin capsules

on a Sartorius XM 1000 P balance for calibration runs and test runs,

usually measuring between 90 and 250 micrograms.  Acetanilide is run as

an ordinary sample (test run) to test the accuracy of the calibration

of the machine each day.  Blanks are combusted 25mm Whatman GF/F

filters soaked in surface sea water and treated in exactly the same

manner. Data are collected and stored by a microcomputer.